## Index
Document ID | Deliverable/ Document | Category of Deliverable | Milestone of Deliverable | Description of Deliverable | Requirement for the Deliverable | Dependecies for Deliverable | Due Date | Status | steps to create | resources needed | owner
| System Architecture Diagram |  | Product Architecture & R&D
| Algorithm Core Development |  | Product Architecture & R&D
| OR Agent Framework |  | Product Architecture & R&D
| Technology Roadmap |  | Product Architecture & R&D
| System Requirements Specification |  | Product Architecture & R&D
| Alpha Prototypes |  | Pre-Partnership/ Post-AAKHS
| Algorithm V&V Report |  | Pre-Partnership/ Post-AAKHS
| Regulatory Strategy Document |  | Pre-Partnership/ Post-AAKHS
| Predicate Device Analysis |  | Pre-Partnership/ Post-AAKHS
| Clinical Performance Targets |  | Pre-Partnership/ Post-AAKHS
| Preliminary Risk Analysis |  | Pre-Partnership/ Post-AAKHS
| Material Selection Rationale |  | Pre-Partnership/ Post-AAKHS
| Design Transfer Package |  | Post-Partnership
| Software Design Controls |  | Post-Partnership
| Manufacturing Work Instructions |  | Post-Partnership
| Test Procedures & Protocols |  | Post-Partnership
| Cybersecurity Framework |  | Post-Partnership
| Contract Manufacturer Assessment |  | Post-Partnership
| QMS Integration Planning |  | Post-Partnership
| Manufacturing Feasibility Assessment |  | Post-Partnership
| Clinical Study Protocol Development |  | Post-Partnership
| Risk Management File |  | QMS
| Software V&V Reports |  | QMS
| Requirements Traceability Matrix |  | QMS
| Design Review Records |  | QMS
| ntegrated QMS Implementation |  | QMS
| Change Control System |  | QMS
| Management Review Records |  | QMS
| Document Control System |  | QMS
| APIs/SDKs Documentation |  | System Integration
| Integration Test Plan |  | System Integration
| Compatibility Matrix |  | System Integration
| Integration Guidelines |  | System Integration
| Platform Integration Execution |  | System Integration
| System Integration Testing |  | System Integration
| Integration Validation Report |  | System Integration
| Final Design Package |  | Design Freeze
| Technology Transfer Agreements |  | Design Freeze
| Contract Manufacturer Qualification |  | Design Freeze
| Process Validation Protocols |  | Design Freeze
| Device Master Record |  | Design Freeze
| Manufacturing Setup |  | Design Freeze
| Algorithm Performance Validation |  | V&V
| Software Validation (IEC 62304) |  | V&V
| Cybersecurity Validation |  | V&V
| Clinical Study Execution |  | V&V
| Biocompatibility Testing |  | V&V
| Cleaning & Sterilization Validation |  | V&V
| IQ/OQ/PQ Execution |  | V&V
| Electrical Safety Testing |  | V&V
| EMC Testing |  | V&V
| System V&V |  | V&V
| SW/AI Technical Sections |  | 510(k) Submission
| Risk Analysis Contributions |  | 510(k) Submission
| Algorithm Documentation |  | 510(k) Submission
| Complete 510(k) Submission Package |  | 510(k) Submission
| Clinical Data Package |  | 510(k) Submission
| FDA Pre-Submission Activities |  | 510(k) Submission
| Regulatory Submission Management |  | 510(k) Submission
| Labeling and Instructions |  | 510(k) Submission
| Algorithm Performance Monitoring |  | Post-Market Surveillance
| Software Updates & Patches |  | Post-Market Surveillance
| Technical Support (Tier 3) |  | Post-Market Surveillance
| Performance Analytics |  | Post-Market Surveillance
| Post-Market Surveillance Plan |  | Post-Market Surveillance
| Customer Support (Tier 1&2) |  | Post-Market Surveillance
| Clinical Data Collection |  | Post-Market Surveillance
| Adverse Event Reporting |  | Post-Market Surveillance
| Field Service & Maintenance |  | Post-Market Surveillance
Index Sheet Summaries
Design Controls & Documentation: This sheet covers all design control procedures required by ISO 13485/21 CFR 820.30. It includes drafting the design and development plan, defining user needs and system requirements, developing the system architecture, and establishing design reviews and verification/validation protocols . Key deliverables (e.g. Design History File, System Architecture Diagram, traceability matrix) are tracked here to ensure all outputs meet regulatory design control requirements.
Risk Management: This sheet implements ISO 14971 risk management. It lists hazard identification, risk analysis (e.g. FMEA), risk/harms assessment, and mitigation strategies. It documents how each identified risk is evaluated and controlled, producing a Risk Management File of hazard analysis tables . The index entry would note that this sheet contains the product’s full risk assessment, risk control measures, and residual risk evaluation.
Software Development & Lifecycle: Focusing on software (IEC 62304), this sheet outlines the entire software life cycle – from the Software Development Plan through coding, unit testing, integration, and system testing. Tasks such as writing the Software Requirements Specification (SRS), defining software architecture, implementing code, and conducting software hazard analysis and V&V are listed. Deliverables like SRS documents, code review records, traceability matrices, and verification/validation reports are tracked. This aligns with IEC 62304’s lifecycle requirements for software safety and traceability .
AI Model Development & Validation: This sheet addresses AI/algorithm tasks (especially if the device includes an AI component). It lists data collection/curation steps, dataset splitting (training/validation/testing), model selection/training, and internal validation (including bias/robustness analysis). It also includes tasks for system integration of the AI model, clinical performance validation (if applicable), and preparation of FDA-required algorithm documentation. The sheet yields an “Algorithm Description” for regulatory submission and documents performance metrics. (No citation available for AI specifics, but it reflects emerging FDA/IEC 62304 guidance on AI transparency.)
Electronics Hardware Development: This sheet covers electronic design and hardware tasks (under IEC 60601-1 requirements). It includes circuit schematics, PCB design, electrical safety analysis, and prototype fabrication. Key steps like component selection, EMI/EMC pre-compliance testing, and regulatory electrical safety verification are listed. Deliverables include electrical schematics, BOMs, and safety test reports. (Index entry notes that this sheet contains all hardware design and safety testing steps.)
Mechanical Design & Engineering: This sheet contains mechanical design tasks. It starts with conceptual design and 3D CAD modeling of the robot components, followed by engineering analyses (FEA, tolerance stacks, load/stress simulation). It includes prototyping (alpha prototypes), mechanical verification tests (accuracy, range of motion, load stability, durability cycles), and design iteration steps. Final deliverables are detailed drawings, DMR (Device Master Record) updates, and test reports. The Index entry would note mechanical design objectives and tests (e.g. device accuracy and stability verification).
Biocompatibility & Material Safety: This sheet implements ISO 10993 biological evaluation. It lists tasks to identify all patient-contact materials, perform a material risk assessment (consider leachables, residues, etc.), and select a test lab. It then details sample preparation and scheduling of ISO 10993 tests (cytotoxicity, sensitization, irritation, etc.). The key deliverable is a Biocompatibility Report. FDA guidance notes that biocompatibility is evaluated on the final device (including after sterilization) . Thus this sheet’s summary would highlight material selection, risk analysis, testing plans, and the final biocompatibility report.
Sterilization Process Validation: This sheet follows ISO 11135 (and related) requirements for sterilization validation. It includes tasks to determine the sterilization method (e.g. EtO, steam) and worst-case load configuration. It documents development of the sterilization validation plan and coordination with a contract sterilizer. The sheet lists IQ/OQ of sterilization equipment and detailed PQ tests (e.g. half-cycle BI tests and full-cycle validations). Final deliverables include sterilization cycle data and a Sterilization Validation Report. As ISO 11135 and FDA process validation guidance explain, sterilization validation comprises IQ, OQ, and PQ stages . The Index entry would reflect these stages and standards.
Packaging & Labeling Validation: This sheet ensures packaging integrity and label compliance. It lists tasks for designing primary/secondary packaging and labels (meeting dimensions, materials, and labeling specs), developing the seal process, and writing a Packaging Validation Protocol. It includes ISTA transport test simulations, seal strength/integrity tests (dye and bubble tests), accelerated aging (shelf-life), and label legibility/adhesion under stress. Deliverables include ISTA test reports, final packaging drawings, and approved label artwork. FDA’s QS Regulation requires labeling to be legible and controlled (MDR-compliant) , so this sheet tracks label content design and verification tests accordingly.
Usability & Human Factors Engineering: This sheet follows IEC 62366/FDA human factors guidance. It includes defining user profiles and use scenarios, task analysis, and identifying critical use-related risks. It then plans and conducts formative usability studies (iterative UI/hardware prototypes with users) and a summative usability (validation) test. Deliverables are the Use Specification, user training materials, and a Usability Engineering Report. (FDA recommends gathering all validation data by the end of design .) The Index would note HF tasks: user interface requirements, formative/summative testing, and IFU review.
Cybersecurity & Data Privacy: This sheet addresses cybersecurity risk management. It lists threat modeling, documentation of security requirements, and implementation of controls (authentication, encryption, network hardening, SBOM creation). It includes cybersecurity testing tasks (vulnerability scanning, penetration and fuzz testing) and remediation of findings. The deliverables are a Cybersecurity Risk Assessment, SBOM, and a Cybersecurity Summary for submission. Under FDA policy, an SBOM listing all software components is required for devices with cyber risk , so this sheet’s summary notes SBOM creation and relevant cyber documentation.
Regulatory Submission & Documentation: This sheet organizes all tasks for preparing a 510(k) (or similar) submission. It starts with developing the regulatory strategy and identifying a predicate device. Key steps include compiling the Device Description, Indications for Use, and Predicate Comparison. It then compiles all test summaries (mechanical, software, electrical safety, EMC, usability, biocompatibility, sterilization, packaging, etc.) into the submission. It includes label and IFU drafts, risk summary, and final assembly of the submission documents (including the 510(k) Summary). The sheet notes that writing and assembling the submission typically occurs in the later months (roughly month 8–9 of development) . The Index entry would reflect that this tab contains the complete premarket submission package with all supporting evidence.
Manufacturing Transfer & Process Validation: This sheet covers the transition to production. It includes creating the Manufacturing Plan and Device Master Record (final release documents), selection and qualification of suppliers/components, and pilot production builds. It details process definition (flow charts, work instructions), and validation of manufacturing equipment/processes: IQ/OQ for tools, and Performance Qualification (PQ) runs. It also covers training production staff and final device verification in the production line. Deliverables include the Process Validation Report and Control Plan. The Index would note process validation for assembly (OQ/PQ) and readiness steps for launch.
Each Index entry above captures the contents of its corresponding tab, providing a clear summary of that deliverable’s scope and standards.

## Design Controls & Documentation
Design Controls & Documentation (ISO 13485 / 21 CFR 820.30) | Unnamed: 1 | Unnamed: 2 | Unnamed: 3 | Unnamed: 4 | Unnamed: 5 | Unnamed: 6
Description & Standards: Establish robust design control processes to ensure the device meets user needs, is safe/effective, and complies with regulatory requirements . Under ISO 13485 and FDA QSR, we must maintain a Design and Development Plan defining stages, reviews, verification/validation, responsibilities, and traceability . Key design control elements include documenting user needs, design inputs (requirements), design outputs (specifications, drawings, code, etc.), design reviews, design verification (does output meet inputs?), design validation (does device meet user needs in intended use), and design transfer to manufacturing . All activities and decisions are recorded in the Design History File (DHF) .
Steps & Tools: |  | Design Controls: Risks, Harms & Mitigations
|  | Task | Risk | Potential Harm (if not addressed) | Mitigation Strategy | Key Deliverables
Design & Development Planning: Draft a comprehensive Design and Development Plan covering all stages of development, planned design reviews, V&V activities, and resource needs as required by ISO 13485:2016 §7.3.2 . Timeline: 1 week by QA/Project Manager; Staff: 1 QA, 1 Project Manager. |  | Requirements gathering | User needs missed or mis-specified; scope creep | Device may not meet clinical needs or require late design changes | Conduct thorough user interviews; include surgeons in reviews; freeze requirements baseline with change control | User Needs Report; System Requirements Specification
|  | Design & architecture | Flawed architecture or interface errors | Integration failures, safety risks (e.g. component incompatibility) | Hold cross-functional design reviews; simulate system behavior (e.g. software-hardware integration tests) early | System Architecture Diagram; Interface Control Doc.
Requirements Definition: Elicit User Needs from surgeons and OR staff (e.g. navigation accuracy, sterilizability, usability needs) and translate them into Design Inputs (technical requirements). Use a requirements management tool (e.g. Jama or DOORS) to capture and link requirements. Timeline: 3–4 weeks with cross-functional input; Staff: Systems Engineer (lead), Clinical/Marketing input, Software/Hardware engineers. |  | Design implementation (outputs) | Design outputs deviate from inputs | Device may not fulfill requirements (performance or safety gaps) | Maintain bi-directional traceability matrix ; internal audits of drawings/code vs. requirements; peer reviews of designs | CAD models & drawings; Schematics; Source code repository
|  | Design reviews & approvals | Inadequate review (issues overlooked) | Design flaws persist, leading to failures or recalls | Use formal review checklists; involve QA, RA, clinical experts in reviews; address all action items before moving forward | Design Review meeting minutes & sign-offs
System Architecture & Specifications: Develop a System Architecture Diagram and allocate requirements to subsystems (mechanical, electronics, software, AI algorithm). Create specifications for each subsystem (mechanical drawings, electrical schematics, software design descriptions). Tools: Onshape for CAD, Altium for PCB design, etc. Timeline: 4–6 weeks; Staff: Systems Engineer, 1 Mechanical Eng, 1 Electrical Eng, 1 Software Architect. |  | Verification testing | Test protocols incomplete or incorrect acceptance criteria | Non-compliance with specs unnoticed; potential safety risks untested | QA review of all test protocols; perform dry-runs of tests; calibrate test equipment (IQ/OQ for test rigs if needed) | Verification Plan; Verification Reports (with trace matrix)
|  | Validation (simulated use) | Validation in unrealistic conditions or with untrained users | Device might not actually meet user needs in real-world (use errors, inefficiencies) | Conduct validation in real OR environment if possible (or high-fidelity simulation); use representative users and scenarios per IEC 62366; train users per IFU then observe performance | Validation Protocol; Validation Report; Use simulation data
Design Reviews: Schedule formal design reviews at key milestones – e.g. requirements review, preliminary design review, critical design review – with cross-functional team (engineering, QA, RA, clinical). Document review minutes, action items, and approvals as required by 21 CFR 820.30(e). Timeline: Ongoing, ~1 week for each review to prepare and follow-up; Staff: Project Manager (organizer), all relevant team leads (R&D, QA, RA). |  | Design transfer to mfg. | Production specs not clear or process not capable | Manufacturing errors, inconsistent device quality (post-market failures) | Early involvement of manufacturing engineering; process capability studies for critical dimensions; pilot production run and process validation (IQ/OQ/PQ – see Manufacturing section) | Device Master Record; Pilot build report; Process Validation docs
|  | Documentation (DHF) | Missing or outdated DHF documents | Regulatory non-compliance; difficulty demonstrating design control in audit
Design Output Development: Conduct detailed design of all system elements: develop 3D CAD models and engineering drawings for mechanical parts, PCB layouts and wiring diagrams for electronics, and code for software/firmware. Ensure outputs meet the input requirements (traceability matrix to be maintained linking each design output to its source requirement ). Use a Product Data Management (PDM) system or QMS repository to control documents. Timeline: 8–12 weeks for initial prototypes; Staff: 2 Mechanical Eng, 1 Industrial Designer, 1 Electrical Eng, 2–3 Software Dev.
Verification Planning: Create a Design Verification Plan detailing how each design input requirement will be verified (e.g. bench test, measurement, analysis). Include equipment needs and acceptance criteria. Timeline: 2 weeks to draft plan; Staff: Systems/Test Engineer, QA.
Design Verification Execution: Perform verification tests on prototypes – e.g. accuracy tests for navigation (position error measurements), electrical safety pre-testing, software unit/integration tests – according to the plan. Document results in verification reports, showing traceability of test results to requirements. Timeline: ~4–6 weeks for execution; Staff: Test Engineer (lead), support from Mech/Elec/Soft engineers for specific tests.
Design Validation: Plan and conduct validation to confirm the device meets user needs and intended use in a real or simulated use environment . This may involve usability studies in an OR-simulation (covered in Usability section) and possibly cadaver labs or simulated surgeries to demonstrate clinical workflow integration. Document a validation protocol and report. Timeline: 4 weeks (after design freeze); Staff: Clinical Engineer or HF specialist (lead), support from engineering.
Design Transfer: Prepare for transfer of design to manufacturing by finalizing the Device Master Record (DMR) – all specifications, drawings, BOMs, assembly procedures – ensuring the production process can consistently produce the device . Conduct pilot builds to verify production instructions. (Manufacturing process validation is detailed in a later section.) Timeline: 4–6 weeks overlapping with late design verification; Staff: Manufacturing Engineer (lead), Mechanical/Electrical Eng.
Documentation & DHF: Throughout, maintain the DHF as a living repository of all design control documents (plans, inputs, outputs, review records, verification/validation plans and reports, risk management file, etc.) demonstrating the design was developed per the approved plan and regulations . Utilize an eQMS software (like Greenlight Guru or similar) for document control and electronic signatures to ensure 21 CFR 11 compliance. Timeline: Ongoing; Staff: QA Documentation Specialist, all document owners.
Timeline & Resources: Design controls span the entire project. Requirements and initial design outputs (drawings, architecture) are targeted for completion in ~3 months (with 4–5 engineers). Verification and validation activities occur later (months 6–9) once a stable design is ready. Staffing: ~1 FTE (full-time equivalent) Systems Engineer for requirements and architecture, plus part-time input from 3–5 engineers (mechanical, electrical, software) during design development, and QA/RA oversight (0.5 FTE each for documentation and reviews). Project Manager ensures schedule adherence and coordinates reviews (0.25 FTE).
Risk Management (Design Controls): A primary risk is inadequate or changing requirements leading to redesign. This is mitigated by thorough upfront user research and traceability – linking outputs and tests back to each requirement to ensure nothing is missed . Another risk is missing a verification step, mitigated by a detailed verification plan and early design reviews. The table below summarizes risks for design control tasks:

## Risk Management
Risk Management (ISO 14971) | Unnamed: 1 | Unnamed: 2 | Unnamed: 3 | Unnamed: 4 | Unnamed: 5 | Unnamed: 6
Description & Standards: Implement a full Risk Management Process per ISO 14971:2019 to systematically identify hazards, estimate and evaluate risks, implement risk controls, and verify risk control effectiveness . All risk activities are documented in a Risk Management File (RMF), demonstrating traceability from each identified hazard through risk analysis, control measures, and residual risk evaluation . This process runs in parallel with design development and is closely linked to design inputs and verification: hazards drive requirements for safety features and testing . We will maintain a Risk Management Plan, Hazard Analysis (e.g. Preliminary Hazard Analysis), FMEAs, and a final Risk Management Report summarizing risk acceptability. The RMF will cover not only device risks (electrical, mechanical, software failures, AI misidentification, etc.) but also use-related risks (addressed via usability engineering) and production/ post-market risks (feedback loops for update). |  | Risk Management: Risks, Harms & Mitigations
|  | Task | Risk | Potential Harm (if not addressed) | Mitigation | Key Deliverables
Steps & Tools: |  | Hazard identification | Some hazards not identified (e.g. rare use scenarios, cyber threats overlooked) | Unmitigated risks could lead to patient harm or recalls post-launch | Use multiple techniques (PHA, brainstorming, use scenario analysis); consult clinical advisors and review FDA databases (MAUDE) for similar device hazards | Hazard List; Preliminary Hazard Analysis document
|  | Risk estimation | Severity/probability mis-estimated (too low) | Inadequate controls for high-risk hazard; patient injury not prevented | Use worst-case assumptions; cross-review estimates in risk review meeting; if uncertainty, treat as higher risk or gather data (e.g. test worst-case scenario) | Risk Matrix with initial risk levels; FMEA worksheets
Risk Management Plan: Create an RM Plan defining the scope of risk management activities, team responsibilities, risk acceptability criteria, and how risk activities will be integrated into development . Include references to applicable standards (ISO 14971, ISO 12100 for machinery if relevant) and how risks will be tracked (e.g. in a risk register database). Timeline: 1 week early in project; Staff: QA/RA lead (Responsible), signed off by PM and management (Accountable). |  | Risk control identification | Chosen mitigations insufficient or introduce new risks | Residual risk remains unacceptable; new hazards from controls (e.g. alarm fatigue) | Apply hierarchy of controls – prioritize design changes over alarms ; do iterative risk reduction until acceptable; perform risk control effect analysis to see if new hazards arise (add to hazard list if so) | Risk Control Implementation Plan; Updated requirements (for safety features)
|  | Implementation of controls | Control not implemented correctly or at all | Hazard not actually mitigated (safety feature fails) | Trace each control to design outputs and verify implementation via testing ; conduct design reviews focusing on safety features; for software, include hazard mitigation requirements in SRS and test cases | Risk Control Traceability Matrix; Verification test protocols/results for each control
Initial Hazard Identification (PHA): Conduct a Preliminary Hazard Analysis to identify potential hazards in the system. Consider all hazard areas: electrical shock, moving robotic parts (mechanical injury), navigation errors leading to surgical injury, software/algorithm failures, cybersecurity threats, use errors, etc. Use brainstorming and past device incident databases. Document hazards, foreseeable sequences of events, hazardous situations, and potential harms (severity). Timeline: 2 weeks for initial analysis; Staff: Systems Engineer + representatives from hardware, software, clinical (brainstorm team). |  | Residual risk evaluation | Residual risk incorrectly judged acceptable | Users exposed to undue risk (e.g. if risk was underestimated or not communicated) | Have management review all “acceptable” residual risks; perform formal benefit-risk analysis for any high residual risks and document justification ; include any warnings in labeling if needed | Risk Management Report; Benefit-Risk Analyses (for any high residual risks); IFU warnings/precautions
|  | Risk file documentation | Poor traceability or incomplete records | Compliance failure (audit nonconformity); difficulty proving risk controls in submission | Maintain a living risk log with hazard-to-harm-to-control traceability ; QA to audit RMF completeness before design freeze; use software tool or matrix to organize links | Risk Management File (comprising Risk Management Plan, Hazard Analysis, FMEAs, Risk Control Table, Verification Cross-reference, Risk Report)
Risk Analysis & Estimation: For each hazard/hazardous situation, estimate risk by severity of harm and probability of occurrence. Use worst-case reasonable scenarios (ISO 14971 definitions). Tools: a Risk Register (spreadsheet or dedicated tool) to record each hazard, hazardous situation, harm, severity, probability. Incorporate FMEA techniques: e.g. Design FMEA for mechanical/electrical failures, Software FMEA for software anomalies, Use FMEA for use errors. Remember that FMEAs analyze failure modes but need mapping to harms to be ISO 14971-compliant . Timeline: 4 weeks to complete analyses iteratively as design matures; Staff: Risk Engineer or QA (facilitator), input from design engineers and clinical. |  | Post-market risk updating | Failure to update risk assessment with field data | New hazards or occurrence rates not handled, leading to incidents or regulatory action | Establish post-market surveillance plan (collect complaint data, user feedback, literature); periodic risk file review (e.g. annually or upon significant change) as required by ISO 14971; have a plan for field safety corrective actions if new risk emerges | Post-Market Surveillance Procedure; Annual Risk Review records; Updated RMF after changes
Risk Evaluation: Compare each identified risk against pre-defined acceptability criteria (from the RM Plan). Determine which risks are unacceptable or ALARP (as low as reasonably possible) and thus require mitigation. Document the evaluation rationale. Timeline: Concurrent with analysis; Staff: Same as above.
Risk Control Measures: Define and implement risk control measures for unacceptable risks. Controls follow the hierarchy: inherent safety by design (e.g. mechanical guards, force limits on robot), then protective measures (hardware or software safety features, alarms), then information for safety (warning in labeling) . Assign each risk a or multiple control measures (e.g. hazard: robot arm collision – control: safety interlocks and motion limit software; hazard: navigation inaccuracy – control: registration verification prompt, etc.). Update design requirements to include these safety features (linking risk controls to design inputs) . Timeline: Ongoing as part of design – begin identifying controls during concept and refine through design reviews; Staff: Design Engineers (implement controls), QA (ensure traced in risk file).
Verification of Controls: Verify and validate that each risk control is implemented correctly and actually reduces the risk. For example, if a control is a software alarm, test that the alarm triggers under the intended fault condition; if a mechanical load limit is set, test that the force never exceeds the safe threshold. Document these verifications (may be part of design V&V test cases, traceable to the risk controls) . Timeline: with design verification (month 6–7); Staff: Test Engineer, QA.
Residual Risk & Benefit-Risk Analysis: After implementing controls, re-estimate the residual risk for each hazardous situation. Ensure all residual risks are acceptable or if not, perform a formal benefit-risk analysis (justification that medical benefit outweighs the remaining risk) and consider further mitigations if needed. Document any risks that require user labeling to control (warnings/precautions in IFU). Timeline: late in development (just before final review); Staff: QA/RA with engineering input.
Risk Management Report: At design completion, compile a report confirming the risk management process was carried out per ISO 14971 and that overall residual risk is acceptable. This report will be part of the 510(k) submission (typically a summary of how risks were mitigated to acceptable levels). Timeline: 1 week at project end; Staff: QA/RA.
Risk Traceability & Updates: Maintain a Risk Traceability Matrix mapping each identified hazard to corresponding risk controls and verification/validation tests . This will ensure completeness and is crucial for audits. Also plan for post-market risk management: establish procedures for monitoring field data (complaints, adverse events) and updating the risk file throughout the product lifecycle . Timeline: Ongoing; Staff: QA (owner), cross-functional team in post-market.
Timeline & Resources: Risk management begins at project kickoff and continues through post-launch. Initial hazard analysis and FMEAs are done early (month 1–2), with iterative updates as design details emerge. Expect to spend 4–6 weeks spread over the design phase to build the risk file, plus periodic reviews (e.g. risk review at each design review). Staffing: Part-time involvement from many team members – led by a Risk Management lead (QA or Systems Engineer) ~0.3 FTE, plus inputs from mechanical, electrical, software, clinical reps during analysis workshops (~0.1 FTE each during those periods). Executive management will be Accountable for risk acceptability decisions .
Risk Management Process Risks: A major risk is incomplete hazard identification (missing a hazard that could cause harm). Mitigation: involve diverse perspectives (engineering, medical, end-users) in hazard analysis and review against standards or databases of similar devices. Another risk is incorrect estimation of risk, mitigated by conservative assumptions and possibly testing to inform probabilities. It’s critical to maintain traceability – failure to trace hazards to controls could lead to compliance issues . The table below outlines specific risks in the risk management process:

## Software Development
Software Development & Lifecycle (IEC 62304) | Unnamed: 1 | Unnamed: 2 | Unnamed: 3 | Unnamed: 4 | Unnamed: 5 | Unnamed: 6 | Unnamed: 7
Description & Standards: The navigation system’s software (including robot control, navigation algorithms, and user interface) will be developed under a rigorous Software Development Life Cycle (SDLC) compliant with IEC 62304. IEC 62304 defines processes for software planning, requirements, design, implementation, verification, risk management, and maintenance . We will classify the software as Safety Class C (since failures could lead to serious injury in surgery, e.g. incorrect guidance) , which imposes the highest level of process rigor. All software work will also align with FDA’s guidance “General Principles of Software Validation” and tie into the overall risk management (software hazards to be addressed). Key deliverables include a Software Development Plan, Software Requirements Specification (SRS), Software Architecture Design, detailed design descriptions (as needed), Hazard Analysis specific to software (linking to system risk analysis), Unit, Integration and System (Verification) Test Protocols and Reports, and a Software Release version for testing and submission. We will use modern tools for configuration management (e.g. Git), issue tracking (e.g. JIRA), and automated testing where possible. |  |  | Software Development: Risks, Harms & Mitigations
|  |  | Task / Area | Risk | Potential Harm (if not addressed) | Mitigation | Key Deliverables
|  |  | Requirements (SRS) | Missing or unclear software requirements (especially around safety or performance) | Critical functionality might be left out or mis-implemented (e.g. no alarm on tracking failure) leading to unsafe operation | Involve end-users and risk analysis in requirements phase; conduct requirements review with cross-functional team; maintain traceability matrix – each risk control must map to a software requirement | Software Requirements Specification; Requirements Traceability Matrix
Steps & Tools: |  |  | Software design & architecture | Poor architecture (e.g. inadequate separation of safety-critical modules, or performance bottlenecks) | Hard-to-maintain code, potential for uncontained failures propagating (a bug in one module crashes entire system) | Apply IEC 62304 design principles; separate safety features into independent modules (e.g. watchdog process); do design reviews with experienced software architects; prototype tricky components to test architecture early | Software Architecture Document; Detailed Design Specs for critical modules
|  |  | Implementation (coding) | Software bugs introduced (logic errors, memory leaks, race conditions) | Device could malfunction: e.g. navigation miscalculations, system freeze during surgery – patient injury | Enforce coding standards and peer reviews; use static code analysis tools to catch common errors; incremental testing (unit tests for each module); simulate high-stress conditions (long runs, high data rates) to find issues | Source Code Repository; Static Analysis Reports; Code Review records; Unit Test Results
Software Development Plan: Draft a plan outlining the software life cycle model (we may use an Agile approach with defined iterations, or V-model aligning with hardware milestones), the development environment, configuration management, and review/approval processes. Include provisions for software risk management and classification. Timeline: 1 week; Staff: Software Lead, QA. |  |  | Software verification | Incomplete test coverage or poorly defined test cases | Some requirements or use scenarios not verified – latent bugs remain | Develop a thorough Software Verification Plan with traceability to all SRS items; include edge cases and stress tests; use test automation to expand coverage (e.g. fuzz testing for inputs) | Software Verification Protocol & Report; Test Traceability Matrix (requirements-to-tests)
|  |  | Schedule for software | Underestimation of software complexity causes delays | Project timeline slips, impacting V&V start; if rushed, quality may suffer (bugs not fixed) | Adopt time-boxed iterations to frequently reassess progress; prioritize critical features first (must-have vs nice-to-have); daily stand-ups to identify blockers; if needed, add temporary contractors for specific tasks (e.g. GUI development) | Iteration Burn-down charts; Progress reports; Adjusted schedules with additional resources if needed
Software Requirements Specification (SRS): Gather and write detailed software requirements based on system requirements allocated to software. This includes functional requirements (e.g. “calculate instrument position within 1 mm accuracy”), performance requirements (latency, real-time constraints), interface requirements, and safety requirements (fault detection, alarms). For any AI components, specify data requirements and performance metrics. Use a requirements management tool to uniquely identify and trace each requirement. Timeline: ~3–4 weeks; Staff: Software Systems Engineer, AI specialist (for algorithm-specific reqs). |  |  | Software-human interface | Poor usability of software (not addressed early) | Use errors by clinicians (e.g. confusing UI could lead to improper use of nav data) | Integrate human factors early: have UX designer work with devs; conduct formative usability tests on software UI (even if simulation); adhere to IEC 62366 (see Usability section) | UI prototypes; Formative Usability Test reports; Updated UI based on feedback
|  |  | Configuration management | Loss of version control or build consistency | Could release a software build with untested changes, or inability to reproduce a version for bug fix | Strict use of version control for all code; define “critical branch” for release; automated build and record version hash; perform configuration audit before release to freeze the code baseline | Software Configuration Index; Release Notes; Version Control logs; Build verification checklist
Software Architecture & Design: Create a Software Architecture Document (SAD) describing the high-level design: software items and components, module decomposition, data flow, state machines, and interfaces. Clearly identify safety-critical modules that require extra scrutiny. Ensure architecture supports risk controls (e.g. independent monitoring tasks for safety). For Class C, document detailed design of key algorithms and modules (Software Detailed Design Specifications) as needed for clarity. Timeline: 3–4 weeks; Staff: Software Architect, with input from electronics (for firmware interfaces) and AI developers. |  |  | Software cybersecurity | Vulnerabilities left in code (e.g. open ports, weak encryption) | Hacking risk: attacker could alter navigation output or disable system, causing patient harm or data breach | Include cybersecurity requirements in SRS (encryption, authentication); run vulnerability scanners on code; do penetration testing on device (see Cybersecurity section) and fix issues; keep SBOM for third-party components to patch known CVEs | Cybersecurity requirements in SRS; Static code scan reports; SBOM; Penetration Test report (issues resolved)
|  |  | Maintenance & updates | Lack of plan for software updates/patches post-launch | Security issues or bug fixes not deployed, leading to accumulative risk in field; potential compliance issues with FDA (524B requirements)
Software Hazard Analysis: As part of risk management, perform a software hazard analysis (required for moderate/high concern software) – identify potential software failure modes or malfunctions that could lead to hazards (e.g. software hangs causing robot to stop or incorrect navigation display causing surgeon misalignment). Link each identified software hazard to mitigation measures such as software requirements for self-tests, watchdog timers, exception handling, or redundant checks . This analysis will often be integrated with the system risk FMEA but also summarized separately for the FDA submission. Timeline: 1–2 weeks; Staff: Software Lead, QA.
Coding & Implementation: Develop the software according to the design. Use coding standards (e.g. MISRA C++ if in C++ or equivalent) to reduce errors. Version control via Git (set up a repository with branching strategy for development vs. release). Perform regular code reviews (peer reviews) for critical code sections (safety-related). Utilize static analysis tools (e.g. SonarQube, Polyspace) to catch bugs and security issues early. Timeline: Ongoing for ~3–6 months in iterative sprints; Staff: Team of ~3–5 software developers (including 1–2 focusing on AI model integration).
Unit Testing (Verification): Write and execute unit tests for individual software modules (especially critical algorithms). Aim for high coverage (e.g. >90% for Class C code). Use a unit test framework (like GoogleTest for C++ or pytest for Python, etc.) and possibly test automation in CI pipeline. Document the Unit Test Protocols and Reports, including pass/fail and any issue remediation. Timeline: Ongoing with implementation – each sprint includes unit tests for new code; Staff: Developers (each responsible for testing their code, reviewed by software test engineer).
Integration Testing: After unit testing, perform integration tests to verify that software components work together (e.g. robot control loop with sensor inputs, or UI with underlying logic). This can include hardware-in-the-loop testing where software is run on the actual electronics or a simulator to ensure real-time performance. Integration test cases will be derived from the SRS and architecture (testing interfaces and data flow). Timeline: ~4 weeks after coding phase; Staff: Software Test Engineer (lead), developers support.
Software Verification (System Testing): Develop a Software Verification Test Plan aligning with IEC 62304 and FDA documentation needs. This includes end-to-end functional tests of the software within the device (as part of overall system verification). Key scenarios: navigation accuracy under various conditions, response to error conditions (network loss, sensor failure), cybersecurity tests (some overlap with dedicated security testing), etc. Execute tests on the integrated system (alpha or beta prototype). Document results in a Software Verification Report, showing traceability to software requirements (each requirement tested or justified). Timeline: ~4 weeks; Staff: Software Test Engineer, Systems Engineer, QA oversight.
Anomaly Tracking & Resolution: Throughout testing, log software bugs in an issue tracker. Use a severity classification for defects (critical, major, minor) and ensure all critical/major issues are resolved or mitigated prior to release. Conduct root cause analysis for high-severity issues and add regression tests if needed. Maintain a revision history of software versions and changes (this will feed into the submission as part of the software description). Timeline: Ongoing; Staff: Software team, QA.
Cybersecurity Implementation: (Detailed in Cybersecurity section, but tightly integrated in software dev) – incorporate security features (authentication, encryption, etc.) during coding, and perform code vulnerability scans. Ensure compliance with Secure Coding practices .
Software Release & Configuration Management: Establish procedures for formal software release (versioning scheme) and configuration management for software items. Perform a final configuration audit to ensure the delivered software version (for verification/validation) matches the specification and all test results are for that version. Archive the build (source code and executable) in the DHF. Timeline: at verification completion; Staff: Software Lead, QA.
Software Validation (IEC 62304 & FDA): Plan a Software Validation (distinct from verification) to ensure the software meets user needs and intended uses in the actual or simulated use environment. This overlaps with system validation and usability testing. For example, validate that the navigation software indeed improves surgical accuracy in a simulated surgery scenario. FDA expects software of moderate concern to be validated via clinical or simulated use if possible. Document the validation activities and results. Timeline: with system validation (month 8–9); Staff: Systems Engineer, Clinical/Usability lead.
Documentation for 510(k): Compile software documentation for regulatory submission: Software Description, Level of Concern rationale (likely “Moderate” or “Major” for AI-based nav system), SRS, architecture diagrams, software hazard analysis, traceability matrix from requirements to test cases, summary of verification/validation results, and Unresolved Anomalies Report (list of any minor bugs remaining). Ensure this package follows FDA guidance and IEC 62304 evidence. Timeline: ~2 weeks to prepare submission docs post-testing; Staff: Software Lead, RA.
Maintenance & Support Planning: Draft a Software Maintenance Plan for post-market (required by IEC 62304) – including how updates will be managed, re-validation needed for patches, and cybersecurity monitoring (tie-in with post-market surveillance). Identify tools for field updates (secure OTA update mechanism, etc.). Timeline: Late in project (just before launch); Staff: Software Lead, Service/Support Engineer.
Timeline & Resources: Software development is one of the longest critical paths. Key milestones: SRS complete by month 2; architecture by month 3; first integrated software prototype by month 5; software verification tests by month 7. Since iterative development is expected, we plan for multiple software builds (e.g. an Alpha software for internal testing by month 5, Beta software for verification by month 7, and Release Candidate for validation by month 8). Staffing: ~3–5 Software Developers (including 1–2 dedicated to AI algorithm coding) over 6–8 months. Additionally, 1 Software Test Engineer from month 4 onwards for test development, and part-time QA for process oversight. The software team works closely with systems engineering and hardware teams for integration.
Risk & Mitigation (Software): Software risks include the possibility of software bugs leading to device malfunction or hazardous situations. Early and thorough testing is the primary mitigation: unit/integration testing and static analysis to catch issues early, and robust verification. Another risk is delays due to rework if requirements change or are missed – mitigated by freezing SRS early and managing changes through a controlled process. Software cybersecurity vulnerabilities are a significant risk; mitigation includes following secure coding practices and performing vulnerability scans/pen tests (see Cybersecurity section). The table below highlights some software-specific project risks:

## AI Model Development & Validati
AI Model Development & Validation | Unnamed: 1 | Unnamed: 2 | Unnamed: 3 | Unnamed: 4 | Unnamed: 5 | Unnamed: 6 | Unnamed: 7
Description & Standards: The system likely includes an AI or machine-learning model (e.g. for image recognition, sensor fusion, or surgical guidance optimization). AI components require special attention to data quality, algorithm bias, and performance consistency. We will manage the AI development as part of the software process but highlight it separately due to additional validation needs. This will follow principles of Good Machine Learning Practice (GMLP) as outlined by FDA/IMDRF . Key tasks include assembling training and test datasets representative of the intended patient population, training the model in a controlled environment, and verifying its performance on independent validation datasets and possibly in simulated clinical scenarios. The AI model must be locked for the 510(k) (any future model updates would require change control or possibly a new submission if impacting safety). We will also mitigate AI-specific risks (like overfitting, dataset bias, cybersecurity of the model). Documentation such as a Algorithm Description Document, Model Training Protocol/Report, and AI Performance Validation Report will be created.
AI Model Development: Risks, Harms & Mitigations
|  |  | Task/Phase | Risk | Potential Harm (if not addressed) | Mitigation | Key Deliverables
Steps & Tools: |  |  | Data quality & representativeness | Training data not representative of all use cases (e.g. missing certain anatomies or patient groups) | Model may perform poorly or fail on certain patient populations – could lead to navigation errors for those cases | Curate diverse dataset covering all intended patient groups ; perform bias checks (model performance by subgroup); if gaps, acquire additional data or include safety margins (e.g. default to safe mode if out-of-distribution input) | Data Management Plan; Dataset Composition Summary; Bias Analysis Report
|  |  | Model accuracy shortfall | AI model does not achieve required accuracy or reliability | Navigation could be erroneous – if undetected, surgeon might be misled (patient injury) | Set clear acceptance criteria from clinical requirements; iterate on model (feature engineering, architecture tuning); if still inadequate, consider hybrid approach (AI suggestions with human verification) or limit use of AI features until improved | Model Performance Metrics Report; retraining logs; Possibly design change if criteria unmet (e.g. improved sensor input)
Data Collection & Curation: Identify and collect data needed to train the AI model (e.g. pre-op images, intra-op tracking data, patient anatomical variations). Ensure datasets are representative of the intended use population – e.g. cover different anatomies, pathologies, sexes, ethnic groups to avoid bias. Label the data as needed (possibly via clinical expert annotation). Establish data agreements or use open datasets if available. Timeline: Ongoing, but initial dataset in first 1–2 months; Staff: Data Engineer, Clinical SME. |  |  | Overfitting/ poor generalization | Model performs well on training data but fails on new data | Unreliable performance in real cases – unpredictable failures | Use separate validation and test sets strictly ; use cross-validation; apply regularization; test on external datasets if possible; integrate a real-time check on confidence (if model is unsure, prompt user or revert to manual mode) | Training & Validation Report; Learning curves; Overfitting analysis (training vs. test performance)
|  |  | Integration issues | Model is too slow or resource-intensive for real-time use | System lag or crash during surgery – unsafe condition | Optimize model (e.g. compress network, use efficient libraries); test throughput on target hardware early; if needed, add hardware acceleration (GPU) or simplify model for deployment | Integration Test Results (latency, throughput); Optimized Model documentation (pruning, quantization if used)
Dataset Split & Management: Split data into training, validation, and test sets – with test set kept strictly independent until final evaluation . Possibly use k-fold cross-validation for training if data limited. Use tools like Python (TensorFlow/PyTorch for ML, pandas for data handling). Ensure data is stored securely and patient information is handled per HIPAA if applicable (data de-identification). Timeline: 1–2 weeks; Staff: AI Engineer (Data Scientist). |  |  | Lack of transparency | Model outputs are not interpretable, and system gives no explanation | Surgeon may not trust or understand AI outputs, possibly misusing it or ignoring important cues | Implement UI to display confidence or rationale (if possible) for AI outputs; include the human-AI team performance focus as per GMLP (design workflow so AI assists, not overrides, clinician); provide training to users in IFU about AI limitations | AI User Interface elements (confidence indicator); User Training materials about AI; Human factors evaluation of AI use
|  |  | Verification & validation | Not enough independent testing (or test data overlaps training data inadvertently) | Overestimation of model performance – latent errors go undetected until clinical use | Maintain strict dataset separation; have third-party or independent team verify results; perform additional real-world simulation tests (e.g. on cadavers or phantom) to confirm real-world performance; incorporate AI into summative usability test to see how users interact with it | AI Verification Report; External Validation Study (if done); Summative test results involving AI decisions
AI Model Development: Select appropriate model architecture (e.g. a neural network) or algorithm for the task. Implement the model and train it on the training dataset. Use best practices to prevent overfitting (regularization, etc.). Track training metrics epoch-by-epoch. Possibly use AutoML or hyperparameter tuning techniques to optimize. Save intermediate models and pick best performing (on validation set) as the candidate. Timeline: Model prototyping by month 3, final model training by month 5; Staff: AI/ML Engineer, with computational resources (maybe utilize cloud GPUs). |  |  | Regulatory compliance | Insufficient documentation of AI (training process, datasets, etc.) for FDA | Submission deficiencies – FDA delay RTA (refuse-to-accept) or additional information requests | Follow FDA’s AI/ML submission guidelines: provide detailed algorithm description, dataset summaries, adherence to guiding principles (e.g. separate training/test sets, multidisciplinary input) ; include Software Bill of Materials covering AI libraries too (for cybersecurity) | Algorithm Description for 510(k); Training Data Summary; Model Files with version control; GMLP Compliance Checklist
AI Verification (Internal Validation): Evaluate the trained model on the independent test set that was not seen during training. Compute performance metrics relevant to safety and effectiveness – e.g. accuracy of navigation predictions, sensitivity/specificity if classification, error distribution, etc. Ensure performance meets or exceeds predefined acceptance criteria (these criteria should tie back to clinical requirements, e.g. “95% of points have <2 mm error”). Perform statistical analysis to verify robustness. Document results in an AI Verification Report. Timeline: ~2 weeks for test & analysis; Staff: AI Engineer, Statistician (if needed for analysis).
Bias and Robustness Analysis: Analyze model performance across different subgroups (e.g. does the model perform equally for different patient demographics or anatomical variations?). Check for any biases. Also test robustness: e.g. introduce slight noise or worst-case scenarios to ensure model doesn’t fail unpredictably. If issues are found, retrain or adjust model accordingly (this may require additional data gathering). Timeline: concurrent with verification; Staff: AI Engineer, Clinical advisor.
Model Integration & Testing in System: Integrate the AI model into the navigation software. Verify it functions correctly in real-time within the system (e.g. latency acceptable, outputs consumed properly by rest of system). Test with simulated input streams or replay real recorded surgical scenarios to see end-to-end system behavior. Timeline: month 5–6; Staff: Software & AI Engineer together.
Clinical Performance Validation (if applicable): Design a study or test (bench or simulated) to validate that the AI’s performance translates to clinical benefit. For example, run the navigation system with AI on cadaveric models or phantom and assess surgical accuracy/outcomes compared to not using AI or compared to expected standard. While a formal clinical trial may not be required for 510(k) if substantial equivalence is claimed, a limited clinical validation or user evaluation can de-risk the AI. Timeline: month 7–8; Staff: Clinical Engineer, possibly external orthopedic surgeon consultants to evaluate results.
FDA Documentation (Algorithm): Prepare an Algorithm Description for the 510(k) that describes the AI’s methodology (e.g. type of model, inputs/outputs, how it was trained), and a summary of performance. Also address the “Predetermined Change Control Plan” if any future model update is anticipated (though for 510(k), usually the model is fixed – but FDA expects discussion if the model can improve with more data). Provide the training dataset summary and describe measures taken to assure generalizability . Possibly include the list of GMLP guiding principles addressed (e.g. data representative, separate test set, human-AI interaction considered) to demonstrate adherence to best practices. Timeline: ~1 week to compile; Staff: RA with input from AI Engineer.
Regulatory/Quality Steps: If the AI will be continuously learning or updated, implement controls to disable live learning (most likely, the model will be static for this regulatory version). Also, include the model in the configuration management (so its exact version is tracked in the DHF). Validate any software tools used in model training that could impact final model (e.g. if using a custom script to preprocess data, ensure it’s verified). Timeline: Ongoing; Staff: QA.
Timeline & Resources: AI development runs in parallel with general software. Data collection/annotation in early phase (months 1–2), model training and internal verification mid-phase (months 3–5), final integration and testing by months 6–7. Staffing: 1 dedicated Machine Learning Engineer/Data Scientist throughout (to develop model), plus partial support from a Clinical Specialist (to ensure clinical relevance of data and output) and Software Engineer (to integrate model). We may leverage external data annotation services for labeling if needed (budget for that).
Risk & Mitigation (AI): Key risks: model performance may be insufficient (e.g. not meeting accuracy needed) – mitigated by iterating on model, enlarging training data, or simplifying the task for reliability. Overfitting risk – mitigated by using separate test sets and possibly external validation; if model doesn’t generalize, consider more data or simpler model. Bias risk – mitigated by ensuring diverse training data and performing subgroup analysis. Lack of transparency in AI decisions – mitigated by choosing interpretable features or adding user alerts when AI uncertainty is high (e.g. flag if confidence below threshold). Regulatory risk: FDA scrutiny of AI is high; mitigated by adhering to GMLP principles and thoroughly validating. The table below addresses AI-specific development risks:

## Electronics Hardware Developmen
Electronics Hardware Development (IEC 60601-1 & IEC 80601-2-77) | Unnamed: 1 | Unnamed: 2 | Unnamed: 3 | Unnamed: 4 | Unnamed: 5 | Unnamed: 6
Description & Standards: The electronic hardware includes the robot’s control electronics, power supply, sensors (e.g. optical trackers or electromagnetic sensors for navigation), and possibly custom PCBs for interface and safety circuits. We will design and verify the hardware to meet medical electrical equipment safety standards IEC 60601-1 (General Requirements for Basic Safety & Essential Performance) and the particular standard IEC 80601-2-77:2019 for Robotically-Assisted Surgical Equipment, which addresses specific safety/performance aspects of surgical robots . Key considerations are electrical safety (isolation, leakage currents), electromagnetic compatibility (EMC per IEC 60601-1-2), and functional safety of control circuits (perhaps using IEC 61508 principles for critical controls). Deliverables include circuit schematics, PCB layouts, BOMs, a Hardware Design Document, engineering test reports (e.g. power consumption, noise), and successful completion of external lab type tests for safety/EMC. We will also conduct hardware risk analysis (component failures, overheating, etc.) and incorporate necessary redundancies or monitoring circuits (e.g. emergency stop).
Electronics Hardware: Risks, Harms & Mitigations
|  | Task/Area | Risk | Potential Harm (if not addressed) | Mitigation | Key Deliverables
Steps & Tools: |  | PCB design & fab | Errors in PCB design (incorrect routing, footprints, etc.) leading to board malfunction | Lost time debugging or need for redesign; critical functions not working (e.g. sensor dead) delaying V&V | Schematic and layout peer reviews; run ERC/DRC checks in CAD; perhaps build one section of circuit on breadboard for proof of concept before PCB; plan two prototype iterations (schedule for re-spin) | PCB Schematics & Layout files; PCB Review checklist; Prototype PCBs (Rev A, Rev B)
|  | Electrical safety | Design not meeting isolation or leakage requirements (e.g. improper grounding, insufficient isolation on power) | Shock hazard to patient or user; device not regulatory compliant – cannot be cleared | Follow 60601-1 guidance from design phase (use isolation components where needed, double insulation on patient connections); get expert consultant review for safety circuits; early dielectric leakage test on prototype using hipot tester in-house | Safety Analysis Report (60601-1 compliance checklist); Test records of leakage current on prototype; Grounding diagram
PCB Design & Schematics: Create detailed schematics for any custom PCBs (e.g. main control board, motor driver boards, sensor interface board). Use ECAD tools like Altium Designer or OrCAD. Critical elements: redundant sensing for critical parameters (if needed for safety), inclusion of watchdog timers, and designing to minimize EMI. Generate PCB layouts following spacing/creepage for mains voltage if applicable and proper grounding. Timeline: Initial schematics by month 2, PCB layout by month 3; Staff: Electrical Engineer (PCB designer). |  | Component failure modes | Key component could fail in dangerous way (e.g. motor driver short causing uncontrolled movement, sensor giving false data) | Patient or user injury (uncontrolled motion, incorrect navigation info) | Implement redundant safety circuits: e.g. current limiting, watchdog that cuts power on fault, independent position sensor to cross-check; perform component FMEA – ensure no single component failure leads to catastrophic harm without detection | Hardware FMEA; Circuit safety features description (e.g. E-stop design); Test results for fail-safe (tripping circuit, etc.)
|  | Thermal management | Overheating of electronics (processor or power components) | Device failure during surgery or burn hazard if user touches overheated part | Thermal analysis (use FEA or empirical testing with thermocouples on hot components in worst-case operating mode); add heat sinks or fans if needed; include temperature sensor on board to trigger alarm or shut down if overheating (risk control) | Thermal Test Report (with worst-case temp vs. limits); Thermal mitigation design (heat sink drawings, fan specs)
Component Selection & Sourcing: Select electronic components (microcontroller/processor, motor drivers, power regulators, etc.) focusing on those with medical-grade reliability where possible. Ensure components have necessary certifications or MTBF data. Identify any components that are new or need additional testing (e.g. custom sensors) and plan for their evaluation. Establish relationships with suppliers for critical components (with second-source if possible). Timeline: month 2; Staff: Electrical Engineer, Supply Chain consult. |  | EMC compliance | Device emits too much EMI or is susceptible to interference (e.g. in OR with other equipment) | Could interfere with other OR equipment or malfunction when exposed to common disturbances (risk to patient safety or OR workflow) | Follow EMC design best practices: filter inputs, shielding cables, use twisted pairs, metal enclosure for controller; pre-test emissions and immunity ; if issues, engage EMC expert or add filters; ensure PCB layout considers RF (ground planes, etc.) | EMC Pre-compliance Test Report; EMC Design Features list; Final EMC Test Report (3rd party)
|  | Schedule (lead times) | Long lead times for electronic components or test equipment (or supply chain issues) | Project delay waiting for parts; if parts unavailable, need redesign with substitutes | Identify long lead components early and order them (even before final design if needed); use dev kits for interim testing if custom boards delayed; have backup distributors; for single-source parts, qualify an alternative in design | Procurement Plan for electronics; Inventory of critical parts by prototype time; Alternate component list
Prototype Fabrication: Order prototype PCBs and assemble them (either in-house if capability exists or via a PCB assembly service). Perform bring-up testing: check voltages, basic functionality of each circuit block. Use lab equipment (oscilloscope, DMM, logic analyzer) to verify signals. If multiple iterations are expected, plan for at least two prototype spins before finalizing. Timeline: month 4 for first prototypes, month 6 for refined prototype; Staff: Electrical Engineer, Technician. |  | Integration with mech/software | Electrical interfaces not aligning (connectors mismatch mech design, or timing of signals not matching software expectations) | Last-minute fixes needed (e.g. adapter cables) or system integration delays; potential reliability issues if temp fixes used | Co-design: mechanical and electrical teams coordinate on enclosure and connector placements (do CAD assembly); perform integration testing as soon as possible with partial systems; maintain an interface control document for pinouts, communication protocols shared by HW/SW teams | Integrated CAD model (boards inside mechanical assembly); Interface Control Document (signals, connectors); Integration Test Report (electronics with motors & software)
|  | External test failures | Failing formal 60601 or 80601 tests (EMC, safety) at lab | Delayed submission timeline (re-test needed); increased costs; if late, could slip launch date | Mitigate by thorough pre-tests and design margin; schedule formal testing with some buffer before submission date; have contingency design mods ready (e.g. if emission too high, know where to add ferrite beads or shielding); utilize lab’s preliminary evaluations (many labs offer pre-testing guidance) | Formal Test Lab Report (pass results); If fail: Problem Report and Corrective Action, followed by Re-test Report
Electrical Safety Design Analysis: Review the design against IEC 60601-1 requirements: sufficient isolation for any patient-applied parts (the navigation system likely is not directly applied to patient except maybe through trackers; if so, ensure isolation of any applied part circuits), fuses or current limiting for fault conditions, proper grounding scheme, and emergency stop circuits for the robot. If using a medical power supply, ensure it’s certified; if internal PSU, design isolation transformers appropriately. Perform a fault tree or FMEA focusing on electrical hazards (shock, fire). Timeline: ongoing, formal review by month 5; Staff: Electrical Engineer, QA.
Bench Testing (Engineering Verification): Conduct internal hardware verification tests on electrical assemblies: measure power consumption, heat dissipation (thermal testing on boards and in enclosure, possibly with thermocouples), battery backup if any, etc. Test all I/Os and sensor readings for accuracy and noise. Simulate faults (e.g. unplug sensor, overvoltage) to confirm protective circuits work. Document these engineering tests. Timeline: month 5–6; Staff: Electrical Engineer, Test Engineer.
Integration with Software: By month 5, integrate electronics with software – e.g. have software running on the actual electronics controlling motors, reading sensors. Verify real-time performance (loop rates, etc.). This will flush out any interface issues (timing, analog signal calibrations, etc.). Timeline: month 5–6; Staff: Electrical & Software team together.
Pre-compliance EMC Testing: To avoid surprises at final testing, perform a pre-compliance EMC test on the hardware around month 6 or 7. This involves using an anechoic chamber or EMC test house to do radiated emission scans, and maybe basic immunity tests, on the prototype. Identify any issues (excess noise at certain frequencies, susceptibility to ESD, etc.) and iterate fixes (add shielding, filtering, grounding improvements) in the design before finalizing. Timeline: ~1 week testing + analysis; Staff: Electrical Engineer.
Design Freeze & Documentation: Finalize the hardware design (schematics, PCB, wiring) by about month 7 after any needed tweaks from testing. Release the drawings and BOM into the configuration management system as part of the DMR. Ensure all components are specified and any manufacturing specs (PCB fab notes, assembly instructions) are documented. Timeline: month 7; Staff: Electrical Engineer, QA (doc control).
External Laboratory Testing (60601 series): Schedule formal testing at a certified lab for IEC 60601-1 (electrical safety: leakage currents, dielectric strength, etc.), IEC 60601-1-2 (EMC emissions and immunity), and IEC 80601-2-77 particular requirements for surgical robots (which may include specifics like safety of motion, accuracy, etc., overlapping with mechanical but also control system tests). Prepare the test samples (should be equivalent to final product) and documentation (list of critical components, circuit diagrams, etc. for the lab). Attend testing if possible to observe. If the device includes any wireless tech, include FCC or wireless compliance testing. Timeline: month 8 for testing (assuming product is near final); Staff: Electrical Engineer (Responsible), QA/RA (to communicate with lab).
Test Reports & Compliance: Receive test reports from lab. If any test failed (common in EMC first pass), implement fixes (e.g. add filters, shielding) and re-test affected parts. Document all results. These reports will be included in the 510(k) or at least referenced (FDA expects EMC and electrical safety testing for electronic medical devices). Timeline: month 8–9; Staff: RA to compile for submission, Engineering to implement fixes if needed.
Hardware-related IQ/OQ: If any custom manufacturing equipment for electronics (e.g. in-house PCB programmer, or conformal coating machine), perform IQ/OQ for that equipment as part of manufacturing validation (see Manufacturing section). Mostly electronics will be assembled by contract manufacturers following IPC standards, but any test jigs need verification. Timeline: month 9; Staff: Manufacturing Engineer.
Timeline & Resources: Electronic hardware development is front-loaded (design by month 3–4, prototypes by month 4–5), with testing and iteration through month 7. External compliance testing in month 8. Staffing: 1–2 Electrical Engineers full-time for design and testing. A PCB layout technician or contractor may be used for board layout if the EE focuses on schematics (0.5 FTE). A Test Engineer or technician is needed part-time for bench tests and managing lab tests (0.5 FTE during months 5–9). Tools/Equipment: Onshape (for integrating electronics in mechanical design like board fitment), Altium for PCB CAD, standard EE lab equipment, access to an EMC pre-test facility.
Risk & Mitigation (Hardware): Key risks include PCB failures or re-spins (which can delay project by weeks due to fabrication lead times). To mitigate, we plan for at least two prototype cycles and use design simulation and reviews (e.g. circuit simulation for critical analog parts, peer review of layouts). Electrical safety risk: components or design not meeting 60601 – mitigated by designing to standard from the start and possibly consulting with a test lab early. EMC risk: often first pass failures – mitigated by pre-testing and by using known-good practices (shield cables, proper grounding). Part availability risk: long lead components could delay builds – mitigated by early procurement of critical parts and having alternates. The table below captures hardware development risks:

## Mechanical Design & Engineering
Mechanical Design & Engineering | Unnamed: 1 | Unnamed: 2 | Unnamed: 3 | Unnamed: 4 | Unnamed: 5 | Unnamed: 6 | Unnamed: 7
Description & Standards: The mechanical system includes the robot’s physical structure (robotic arm or navigation device housing), any attachable instruments or markers, and the mounting system (e.g. base or OR table attachments). Mechanical design must ensure structural integrity, accuracy of movement or tracking, and user safety (no pinch points, smooth surfaces, biocompatible materials where needed). It will follow best practices and relevant standards: for example, we consider IEC 60601-1 for mechanical hazards (moving parts, stability), and ISO 14738 for ergonomic design of machinery (to ensure user comfort), as well as applying risk control measures for mechanical hazards from ISO 14971. If specific robotics standards (like ISO 10218 for industrial robots or IEC 80601-2-77 which also covers mechanical safety of surgical robots) apply, we will design to those. Key deliverables: 3D CAD models and engineering drawings of all parts, tolerance analyses, a Mechanical BOM, prototype units, and test reports (e.g. range-of-motion verification, accuracy tests, load tests). We’ll also plan for sterilization compatibility of parts that contact sterile field (e.g. using materials that can withstand autoclave if reusable) – overlapping with Biocompatibility and Sterilization sections.
Mechanical Design: Risks, Harms & Mitigations
|  |  | Task/Aspect | Risk | Potential Harm (if not addressed) | Mitigation | Key Deliverables
Steps & Tools: |  |  | Structural integrity | Underspecified strength – parts could deform or break under load (e.g. arm joint fails) | Device collapse or unintended motion causing patient or user injury; downtime if breaks in surgery | Use conservative safety factors in FEA; choose high-strength materials for load-bearing parts; include mechanical stops to prevent excessive stress; physically load-test prototypes to validate calculations | FEA Reports; Mechanical Test Report (load tests); Updated design with reinforcements if needed
|  |  | Precision & accuracy | Mechanical play or tolerance stack-up leads to positioning inaccuracy beyond spec | Navigation error – could misguide surgeon, risking incorrect cuts or implant placement | Tighten tolerances on critical dimensions; use high-precision components (e.g. low-backlash gearboxes, high-resolution encoders); incorporate a calibration routine to correct minor offsets; test accuracy early and adjust design (e.g. add stiffness) if needed | Tolerance Analysis; Calibration Procedure; Accuracy Test Report (showing system meets requirement)
Conceptual Mechanical Design: Based on system requirements, define the mechanical architecture – e.g. degrees of freedom of the robotic arm, mechanism types (linkages, linear rails, etc.), overall size envelope to fit in OR space. Create rough CAD sketches or physical mockups (perhaps using foam or 3D-printed pieces) to evaluate form factor and surgeon interaction. Timeline: month 1–2; Staff: Mechanical Engineer (Lead), with input from clinical advisor for usability. |  |  | Moving parts hazard | Pinch or crush points present in moving linkage; robot could hit user or patient if uncontrolled | Operator or patient could get pinched, leading to injury (fingers caught, etc.) | Identify all moving link interfaces in risk analysis; cover pinch points with guards or shrouds; implement software motion limits and an E-stop (tie-in with hardware and software); ensure motion speed and force are limited to safe ranges per risk assessment | Risk Control Implementation (guards, covers in CAD; E-stop button on device); Safety Test (simulate obstruction – verify system stops)
|  |  | Usability & ergonomics | Device awkward to set up or operate (too heavy, hard to position) | Incorrect usage or setup errors (e.g. misalignment) leading to prolonged surgery or user fatigue; potential misuse causing hazard | Conduct formative evaluations with clinicians (let them handle prototypes to give feedback); add ergonomic features: handles, balanced weight distribution, quick-release clamps; limit weight of any component to manageable load or provide lift assist if heavy | Formative Study Reports (with feedback and changes); Ergonomic design features (documented in drawings: e.g. handle positions)
3D CAD Modeling: Develop detailed 3D models of each component using Onshape (cloud CAD) or SolidWorks. This includes the robot arm links/joints, the navigation camera or sensor mounts, electronic enclosures, cabling routing, and any disposable or reusable surgical attachments (like tracking markers or bone pins). Ensure clearance through full range of motion, and design out pinch or trap points (or add guards). Use CAD assembly simulations to verify ranges and that the robot can reach required positions. Timeline: month 2–4 (iterative); Staff: 2 Mechanical Engineers. |  |  | Material incompatibility | Chosen materials not compatible with sterilization or not biocompatible (for parts contacting patient) | Sterilization could damage parts (crack, deform), leading to failure; or patient could have adverse tissue reaction to material | Consult ISO 10993 and sterilization method specs: use medical-grade materials (e.g. autoclavable plastics for reusable parts); do biocompatibility testing on any new material coatings; possibly do a few sterilization cycles on prototype to see if any degradation (thermal or chemical) | Material Specification List (with biocomp status); Sterilization Cycle Test on materials; Biocompatibility Test Reports
|  |  | Manufacturing feasibility | Complex part designs or tight tolerance might be very costly or impossible to make at scale | Production delays or cost overruns; quality issues in manufacturing leading to rejects or recalls | Perform DFM/DFA review: simplify part geometry (e.g. split into two simpler parts if needed), loosen non-critical tolerances; engage a manufacturer early (get feedback or prototyping quotes); consider alternate manufacturing methods if needed (e.g. 3D print vs machine vs mold) | DFM Review Minutes; Revised Drawings (post-DFM improvements); Supplier quote feedback incorporated
Material Selection: Choose materials for each part considering strength, weight, and biocompatibility. For patient-contacting or sterile field parts (e.g. instrument trackers), likely use stainless steel, anodized aluminum, or medical-grade plastics (PEEK, Ultem, etc.) that can be sterilized. Ensure any coating or finish is biocompatible (confer with ISO 10993 requirements – see Biocomp section). Timeline: by design freeze; Staff: Mechanical Eng, Materials specialist (consult if needed). |  |  | Environmental factors | Device not designed for OR environmental factors (fluids, cleaning, etc.) | Fluids (blood, saline) ingress could damage device or cause malfunction; cleaning agents could corrode materials; OR lights/temps could affect sensors | Design enclosures to at least IPx1 or IPx2 drip-proof if near surgical field; use seals or conformal coating for electronics if needed; choose corrosion-resistant finishes; verify device functionality after exposure to typical cleaning/disinfection (test per IEC 60601-1 cleaning procedures) | Enclosure sealing design (gaskets, IP rating); Cleaning Instruction Validation (material compatibility test with hospital disinfectants); Post-cleaning test report
|  |  | Integration mechanical-electrical | Mounting points or cable routing improperly planned (e.g. cables too short, or interfere with motion) | Assembly issues (cannot connect components easily) or cable snags that could lead to failure or hazard (cable pulling on something) | Coordinate mechanical and electrical CAD: route cables in CAD models, leave slack and securing points; do a trial assembly with full wiring on prototype to identify issues; provide strain reliefs and cable guides in design | Integrated Assembly CAD (with cables modeled); Assembly Instruction with cable routing; Issue log from trial assembly resolved in final design
Analysis (FEA & Tolerance): Perform engineering analyses: Finite Element Analysis (FEA) on critical load-bearing parts to ensure they can handle worst-case forces (e.g. someone inadvertently leaning on the robot, or static load of attachments). Check deformation to ensure navigation accuracy won’t be compromised by flex. Perform tolerance analysis on assembly stacks to ensure the robot can be assembled consistently and achieve necessary precision. If the device uses precision guides or calibration jigs, ensure their tolerances meet the required accuracy. Tools: FEA software (ANSYS or SolidWorks Simulation). Timeline: month 4–5; Staff: Mechanical Engineer specialized in analysis.
Prototype Fabrication: Build Alpha prototypes of mechanical components. This could involve rapid prototyping: 3D printing for quick form/fit testing of enclosures or jigs, machined parts for structural elements, possibly off-the-shelf components (like linear actuators or standard robot joints) integrated as placeholders. Assemble the prototype in-house. Use these prototypes to verify basic fit, ergonomics (e.g. how a surgeon interacts with it), and to allow software integration testing. Timeline: initial prototype by month 4; Staff: Mechanical Eng, Technician.
Mechanical Verification Tests: Plan and conduct tests to verify mechanical requirements:
Accuracy Test: If the robot positions an instrument, measure the actual vs. expected position (using a high-precision tracking system or laser measurement). Ensure it meets spec (e.g. ≤1 mm error).
Range of Motion: Verify the robot can cover the required workspace (use CAD or physical test, moving to extreme positions).
Load and Stability: Test that the system remains stable (no tipping) under worst-case loads or forces (IEC 60601-1 tilt stability test if applicable). Also test joints under load for any slipping or excessive deflection.
Environmental Tests: If required, check performance under temperature/humidity extremes or after vibration (for shipping). This can overlap with packaging tests (drop tests).
Durability: Cycle test moving parts to simulate usage over time (e.g. move arm back-and-forth thousands of times) to catch any wear issues.
Document all these in a Mechanical Verification Plan & Report. Timeline: month 6–7; Staff: Mechanical Test Engineer, Lab Tech.
Design Iterations: Based on prototype testing, refine the design (Beta prototype). E.g., if a part was too weak or hard to use, update CAD and produce a new version. Conduct a Critical Design Review before freezing design, to approve final mechanical drawings. Timeline: month 5–6 for Beta updates; Staff: Mechanical team.
Finalize Drawings & DMR: Release final 2D engineering drawings with dimensions and tolerances for all custom parts. Include any assembly drawings or instructions. These go into the Device Master Record. Ensure proper version control – drawings signed by engineering and checked by QA. Timeline: month 7; Staff: Mechanical Eng, QA.
Tooling & Manufacturing Prep: If high-volume manufacturing or complex parts (like injection molded plastic covers) are planned, initiate tooling design (injection mold design, casting mold, etc.) after design freeze. Validate tooling via sample runs (this crosses into manufacturing transfer – see that section). For now, ensure manufacturability: design for machining or 3D print for prototypes, design for molding/casting for final if needed. Timeline: month 7 onward (overlap with process validation); Staff: Manufacturing Engineer, Mechanical Eng.
Mechanical Safety and Compliance: Verify that the mechanical aspects meet any applicable standards: e.g., sharp edges test (no sharp edges that could cut – per 60601), mechanical strength (drop test from a certain height if required), ingress protection (if applicable to electronics housing), and any requirements from IEC 80601-2-77 regarding robot mechanical safety (for instance, redundant brake systems or manual override capability). Much of this is covered via risk analysis and verification tests above. Timeline: by end of verification; Staff: QA, Mechanical Eng.
Timeline & Resources: Mechanical design runs from concept in month 1 to design freeze by month 7. Overlapping phases: initial concept (1–2), detailed CAD (2–4), prototypes (4–5), testing (6–7). Staffing: 2 Mechanical Engineers full-time for design. A Mechanical Technician or prototyping shop is needed for fabrication and assembly (on call as needed). If special analyses (FEA) are complex, may use an analysis specialist (consultant) briefly. Tools: Onshape for CAD modeling (allows collaboration and cloud access), possibly SolidWorks or CREO if team is familiar. FEA software for stress analysis. 3D printers (SLA/FDM) for quick prototyping small parts; access to a machine shop (CNC) for metal parts.
Risk & Mitigation (Mechanical): Risks include design not meeting performance (e.g. robot not as accurate or stable as required). Mitigation: iterative prototyping and testing with adjustments; incorporate precision alignment features (dowel pins, calibration routines). User safety risk: moving parts could injure – mitigated by mechanical stops, force limits in software, and training (plus emergency stop accessible). Ergonomic risk: device hard to use or position – mitigated via formative evaluations with surgeons (adjust design accordingly, e.g. add handles or adjust weight distribution). Manufacturability risk: complex parts or tight tolerances could be hard to produce – mitigated by consulting manufacturing experts early and adjusting design (DFM/DFA reviews). The table below covers mechanical design risks:

## Biocompatibiliity & Material Sa
Biocompatibility & Material Safety (ISO 10993) | Unnamed: 1 | Unnamed: 2 | Unnamed: 3 | Unnamed: 4 | Unnamed: 5 | Unnamed: 6
Description & Standards: Any component of the system that contacts the patient directly or indirectly (e.g. surgical instruments or attachments, tracking markers that are fixed in bone or on skin, etc.) must be assessed for biocompatibility per ISO 10993-1 and FDA guidance (Use of ISO 10993-1 within a risk management process). The surgical robotic navigation system likely has invasive accessories (bone pins, reflective markers, etc.) and possibly patient-contact surfaces (e.g. a registration probe, or a drape that contacts patient skin). We will identify the contact classification (e.g. surface-contacting, tissue/bone-contacting, blood-contacting; duration <24h short-term or longer if multiple uses) and determine the necessary biological effect tests: cytotoxicity, sensitization, irritation at minimum for any patient contact, plus systemic toxicity, pyrogen, or sub-chronic tests if indicated by risk assessment . We will create a Biological Evaluation Plan (BEP) that outlines how we’ll demonstrate biocompatibility – either through direct testing of the final device (preferred, using GLP labs) or rationale for leveraging existing data if materials are well-known (with certificates). After testing, a Biocompatibility Report will summarize results and conclusions. Testing will be outsourced to a certified lab (e.g. Nelson Labs or similar) experienced in ISO 10993. All testing will be done on final finished devices (or worst-case equivalent) including any manufacturing residues, as required .
Biocompatibility: Risks, Harms & Mitigations
|  | Issue/Risk | Risk Description | Potential Harm (if not addressed) | Mitigation Strategy | Key Deliverables
Steps & Tools: |  | Material fails biocomp test | e.g. polymer coating causes cytotoxicity or allergic reaction in sensitization test | If not caught, patient could experience tissue irritation, toxicity, or allergic reaction; regulatory rejection if tests fail | Pre-screen materials: request biocomp data from suppliers for plastics/coatings; perform informal cytotoxicity on early sample; use only ISO 10993-tested materials if possible; if test fails, pivot to alternative material or improve cleaning (document rationale and retest) | Biological Evaluation Plan (with material info); Cytotoxicity test results (pass or investigation if fail); Updated material spec if change needed
|  | Test scheduling delays | External lab availability causes test to finish late, compressing submission timeline | Could delay submission filing; rushing may lead to incomplete reports or missing data to FDA | Contact labs early, reserve slots especially for long tests (sensitization); if timeline critical, use multiple labs (split tests) or pay for rush service; closely follow up on progress; have BEP approved internally early to avoid re-planning | Lab SOW/Quote; Testing Schedule; Expedited testing agreement (if used)
Identify Contact Materials & Classification: Catalog all device parts and identify which (if any) have direct or indirect body contact. For example: bone fixation pins (contact bone, implanted during procedure – temporary <24h), tracking markers (contact skin or surgical wound), any robotic end-effector if it touches patient. Determine contact type (surface vs. external communicating vs. implant) and duration (likely short-term ≤24h for each surgery). Timeline: Early, by design phase when materials chosen; Staff: Materials Engineer or RA, with Mechanical Eng. |  | Wrong sample preparation | If samples are not prepared per standard (wrong extraction solvent or time, not sterilized when final device will be sterilized) | Test results could be invalid or not accepted by FDA, causing redo and delay | Strictly follow ISO 10993-12: e.g. use worst-case extraction (surface area/volume ratios, appropriate solvents for polar/non-polar); sterilize samples using final method before testing; consult lab on preparation details; document all preparation steps in reports | Sample Preparation Protocol; Confirmation from lab on receipt of proper samples; Description of sterilization and extraction in Biocomp Report
|  | Over-testing or under-testing | Doing unnecessary tests (wasting time) or omitting required tests | If over-test: lost time/resources; If under-test: FDA may not accept, leading to deficiency letters or need for post-submission testing, delaying clearance | Use risk-based approach per ISO 10993-1 and FDA guidance matrix to determine exactly which tests needed ; justify any omissions with scientific rationale in BEP; if any doubt, consider FDA pre-submission to confirm biocomp plan | Biological Evaluation Plan (with test rationale); Possibly FDA feedback on plan; Completed required tests only (no omissions without justification)
Risk Assessment of Materials: For each material (metal alloy, polymer, coating, adhesive, etc.), see if it has prior usage in similar devices. If using only well-characterized materials (e.g. ASTM F138 stainless steel for bone pins), we might leverage existing biocomp data. However, any novel material or coating must be tested. Use ISO 10993-1:2018 flowchart to determine needed tests based on contact category. For instance, bone-contacting short term likely needs cytotoxicity, sensitization, irritation, acute systemic toxicity, possibly material-mediated pyrogen. If parts are externally communicating to tissue (like something in a surgical wound), similar battery of tests. Document this in the BEP. Timeline: BEP drafting ~2 weeks; Staff: RA or Biocompatibility Specialist. |  | Biocomp documentation gap | Not properly documenting the process or rationale (e.g. material changes not tracked, or no clear link in RMF) | Audit issue or FDA question (they might ask for clarification or extra info) | Keep detailed records: material certificates, lot numbers of tested samples, lab accreditation; integrate biocomp into risk management file (hazard of toxicity marked as addressed by testing) ; review FDA guidance checklist to ensure all needed info (like justification of worst-case article tested) is in final report | Material Certs in DHF; Test Lab accreditation cert; Risk management entries linking to biocomp; Final Biocompatibility Report (comprehensive and formatted for FDA)
|  | Post-market reactions | Rare case: patient reactions after launch (e.g. unforeseen allergy) | Patient harm (e.g. skin irritation in broader population, or allergy to a component) and potential recall or field action | Mitigate by thorough pre-market testing and literature review; include in labeling any known allergens (e.g. if device contains latex or certain metals, though we’d avoid those); set up post-market surveillance to capture any adverse reactions and respond quickly (could involve further testing or material changes) | Labeling (allergen info if applicable); Post-market surveillance plan focusing on biocomp issues; Procedure to initiate CAPA if any biocomp-related complaint arises
Select Test Lab: Engage a lab (such as Nelson Labs, NAMSA, or Toxikon) with our BEP to get their input and quotes. Ensure the lab is GLP-compliant for biocompatibility tests. Schedule testing, keeping in mind some tests (sensitization, systemic toxicity) can take several weeks. Timeline: Start communication by month 5 (once materials and processes are final or nearly final); Staff: RA (Responsible for lab coordination).
Sample Preparation: Manufacture test samples of the parts needing biocomp testing using final production process (including sterilization if the device will be sterilized before use, since residues can affect biocomp). If device is too large to test whole, lab may allow representative portions or extracts. Follow ISO 10993-12 for sample preparation and extraction conditions . Timeline: produce samples by month 6; Staff: Mechanical Eng/Technician to produce, RA to oversee compliance.
Conduct Biocompatibility Tests: The lab will perform tests like:
ISO 10993-5 Cytotoxicity: incubate cell cultures with device extracts, look for cell lysis.
ISO 10993-10 Sensitization: Guinea pig maximization or equivalent test to check if materials induce allergic reaction.
ISO 10993-10 Irritation: e.g. intracutaneous injection test to see if extracts cause irritation in tissue.
ISO 10993-11 Acute Systemic Toxicity: inject extracts into mice to see acute illness.
ISO 10993-11 Material Mediated Pyrogenicity: if applicable, rabbit pyrogen test to ensure no fever response.
If any polymers in long contact, possibly ISO 10993-3 genotoxicity or ISO 10993-6 implantation tests could be needed, but likely not for short-term unless high concern.
The test selection follows our plan and FDA guidance matrix. Timeline: Testing itself ~4–8 weeks depending on test durations (sensitization often longest ~~, pyrogen a few days, etc.); Staff: External lab (executes), RA monitors.
Test Reports & Review: Receive test reports from the lab for each test. Review results: ideally all tests pass (no cytotoxicity or only mild reactivity acceptable per standard, no positive sensitization, etc.). If a test fails (e.g. cytotoxicity is common to fail initially), evaluate cause – sometimes rinsing or additional processing can remove a contaminant. Decide on corrective actions: e.g. change material or improve cleaning if needed, then re-test. Timeline: month 7 for results; Staff: RA/QA review, consult engineers on failures.
Biological Evaluation Report: Compile a report that ties everything together. Include the device description and materials, contact categorization, summary of tests performed and results, and a conclusion that the device is biocompatible for its intended use. If any tests were not done because of rationale (e.g. using same material as a predicate device), document the justification . Ensure this report clearly references ISO 10993 and FDA guidance criteria. This will be included in the regulatory submission. Timeline: month 8; Staff: RA (primary author), reviewed by QA.
Chemical Characterization (if needed): In some cases, doing a chemical characterization (ISO 10993-18) can help identify leachable substances and possibly reduce biological testing. If our risk assessment suggests it, we might send samples for extractables/leachables analysis (GC-MS, etc.) to identify chemicals that come off the device. If those chemicals are below toxic thresholds, that could justify not doing some animal tests. We’ll consider this approach if applicable. Timeline: parallel to testing if chosen; Staff: RA/Chemist.
Ensure Ongoing Compliance: Add biocompatibility to risk management – i.e. the risk of material incompatibility. The RMF should note that biocomp was demonstrated and any risks (like a known sensitization potential of a certain polymer) are mitigated by testing and/or material choice. Also set controls for manufacturing: ensure no unexpected chemical residues (like cleaning agents) remain – this ties into validating cleaning processes in manufacturing or requiring certificates from suppliers. Timeline: Ongoing; Staff: QA, Manufacturing.
Timeline & Resources: Biocompatibility planning (BEP) around month 4–5 when design materials are finalized. Testing in month 6–7, results by month 7–8, with report by month 8 in time for submission compilation. Staffing: Part-time Regulatory Affairs or Biocompatibility Specialist (0.2 FTE for plan and report). External resources: Biocompatibility test lab (cost and schedule to be managed). Tools: Primarily documentation and external lab – internally, ensure we have final product samples (including any sterilization) ready for the lab.
Risk & Mitigation (Biocompatibility): Risks: Test failures – e.g. if material fails cytotoxicity or causes irritation. Mitigation: choose medical-grade materials with known biocompatibility history wherever possible; do an early screening test (like cytotoxicity on prototype materials) to catch issues when changes are still easy. If a test fails, mitigate by material change or additional processing (e.g. extended rinse to remove irritants) and re-test. Another risk is delays in testing (labs can have backlog). Mitigation: schedule in advance and possibly use expedited service if timeline is tight. Also ensure correct sample preparation – e.g. device must be in final finished form (including sterilization) , so coordinate with sterilization timeline. If sterilization validation isn’t done yet, we might sterilize samples with a half-cycle just for biocomp. The table below outlines biocompatibility-related risks:

## Sterilization Process Validatio
Sterilization Process Validation | Unnamed: 1 | Unnamed: 2 | Unnamed: 3 | Unnamed: 4 | Unnamed: 5 | Unnamed: 6
Description & Standards: The system may include components that must be sterile when used (e.g. surgical instruments or disposable attachments that contact tissue). We will identify those components and choose an appropriate sterilization method. For many reusable surgical tools, steam sterilization (autoclave) is common (performed by the hospital), in which case our role is to validate that the device can withstand autoclave and provide validated instructions for use (IFU) for sterilization. If we supply any components sterile single-use, we likely will use Ethylene Oxide (EO) sterilization or Gamma irradiation and must validate that process per ISO 11135 (EO) or ISO 11137 (gamma). For planning, let’s assume we have some disposable parts (e.g. single-use bone pins or drapes) that we will EO sterilize prior to distribution, and possibly some reusable instruments that hospitals will autoclave (so we validate our recommended autoclave cycle under ISO 17665). We will perform a Sterilization Validation including IQ/OQ/PQ of the sterilization process . Likely we’ll use a contract sterilization provider for EO; our role is to develop the validation protocol (half-cycle overkill method, etc.), prepare product loads, and analyze results. We also must ensure the product meets Sterility Assurance Level (SAL) 10^-6 and that residual EO or other byproducts are within safe limits (per ISO 10993-7). Deliverables: Sterilization Validation Protocol & Report, Sterilizer IQ/OQ records (from vendor), BI (Biological Indicator) test results, residual analysis report, and if applicable, a Reusable Device Reprocessing Validation (cleaning + steam sterilization instructions validation per FDA guidance and AAMI TIR12/ISO 17664). IQ/OQ/PQ for sterilization equipment ensures the process is properly installed, operating within specs, and produces sterile product consistently .
Sterilization Validation: Risks, Harms & Mitigations
|  | Risk | Description | Potential Harm (if not addressed) | Mitigation | Key Deliverables
Steps & Tools: |  | Sterilization cycle ineffective | Chosen cycle doesn’t fully sterilize worst-case locations (BI positives in validation) | Non-sterile product could lead to infection in patients; validation failure delays product | Use overkill method (half-cycle) with ample margin ; pre-evaluate packaging design to ensure EO penetration (e.g. use tray with vents, etc.); if failure, adjust cycle parameters (increase gas exposure, etc.) and repeat PQ until success | Sterilization Validation Protocol; BI Placement diagram; PQ BI Results (showing all pass after any needed adjustments)
|  | EO residuals high | EO gas or by-products remain in product above limits (especially if product has absorbent materials) | Toxic reactions in patients (hemolysis, irritation) or regulatory non-approval | Minimize absorbent packaging; add extra aeration time; test residuals on sample – if high, implement additional degassing (warm aeration days) until within ISO 10993-7 limits; consider switching to another method (e.g. gamma) if EO not suitable | EO Residual Test Report; if initial fail: Residual Reduction Action (e.g. extended aeration) and re-test results; Final verification that residual < limits
Determine Sterilization Needs: From the design, list which components need to be sterile. For disposables that we will package and sterilize, decide on method (EO gas is common for mixed materials and because it’s a low-temperature process safe for plastics). For any reusable part, define a reprocessing procedure (cleaning + autoclaving). Include this in risk management (contamination control hazards). Timeline: by design freeze; Staff: RA, Mechanical Eng, Clinical input. |  | Autoclave impact on device | Reusable parts deform or corrode after repeated steam cycles (not initially apparent) | Device failure in field (e.g. part breaks after a few uses) or loss of function (joints stiffen, etc.), causing risk to patient or additional cost | Select materials known for autoclave durability (e.g. PPSU plastic, stainless steel grades); do accelerated life test: e.g. 10× sterilization cycles and inspect; apply protective coatings as needed (anodize aluminum to avoid corrosion); provide protective cases for autoclaving if needed (to avoid mechanical damage) | Reprocessing Validation Report (showing device tolerates X cycles); Material compatibility notes in Biocomp Report for heat/steam; Updated IFU with any limitations (e.g. max cycles or specific autoclave parameters allowed)
|  | Coordination with sterilization provider | Scheduling or communication issues with contract sterilizer (delays, miscommunication of load config) | Delayed validation completion; potential need to repeat if protocol not followed correctly | Start engagement early, clearly define responsibilities; send company representative (or at least be in close contact) during first cycle; double-check BI placement and count with them; have a backup provider if one cannot meet timeline | Supplier Quality Agreement or Validation Agreement with sterilizer; Meeting minutes/ emails confirming cycle details; Contingency plan documented (e.g. alternate site)
Develop Sterilization Validation Plan: If using EO sterilization: follow ISO 11135. Plan for IQ (installation qualification of sterilizer) – likely handled by contract provider but we’ll need their IQ report, OQ – establishing cycle parameters (we might use standard cycle from vendor but ensure our product can handle it), and PQ – the microbiological validation. For EO PQ we will use the half-cycle overkill method : run half the full exposure time with biological indicators (BIs) in product to ensure they are all killed, repeated for 3 half-cycles for sterility assurance . Also plan separate full-cycle runs for Physical PQ to ensure no product damage and to measure EO residuals. If gamma: plan dose mapping. For steam (user reprocessing): plan validation of cleaning and sterilization per ISO 17664 – including worst-case soil testing and microbial kill in autoclave. Document all this in a protocol. Timeline: draft by month 5; Staff: QA/Validation Engineer, RA. |  | Packaging/sterilization interplay | Packaging system fails during sterilization (pouches burst, seals not breathable enough for EO) | Sterilization ineffective or packaging integrity compromised (sterile barrier fail) – infection risk | Work closely with packaging engineers: use packaging validated for chosen sterilization (e.g. Tyvek for EO gas permeation); include representative packaging in validation cycles; do bubble leak tests on packages post-sterilization to ensure seal integrity (tie into packaging validation) | Packaging Design Specs (material, venting for EO); Package integrity test results post-sterilization; Part of Sterilization Report referencing package lot and type used
|  | Documentation incomplete | Not capturing all required data (e.g. missing BI lot info, not reporting worst-case configuration) | FDA may reject sterilization info, requiring additional submissions or testing | Follow ISO 11135 and FDA guidance checklists for sterilization validation documentation; ensure sterilization report states SAL achieved, method, reference to standards, load config, BI results, residuals, etc.; include EO sterilization certificate or compliance letter from vendor for their cycle | Sterilization Validation Report (comprehensive); Test records (BI certificates, BI lot #, residual raw data); Section in 510(k) submission detailing sterilization method, SAL, EO residuals compliance
Coordinate with Contract Sterilizer: Engage an EO sterilization service (e.g. Steris AST). Provide them with product info (dimensions, materials, packaging) to design the validation. Ensure they will perform EO residuals analysis as well. Schedule the validation runs – these often take a few weeks for incubation of BI. For hospital steam, plan an internal or third-party lab validation. Timeline: month 6; Staff: RA/QA.
Sterilizer IQ & OQ: Obtain IQ documentation from contract sterilizer (their chamber calibration, ventilation, etc.). OQ typically involves empty chamber studies to set EO cycle parameters (temp, humidity, gas concentration, time) that achieve desired lethality. If we use their standard cycle, OQ might already be done – we ensure it’s appropriate for our product (no excessive cycles that degrade product). Timeline: provided by vendor; Staff: Vendor primarily, QA to review.
Load Preparation for PQ: Determine worst-case product load for sterilization – typically the maximum load (to ensure EO penetration) or product with the most challenging features for gas penetration. Place Biological Indicators (BIs) (strips with >10^6 resistant spores) inside product at hardest-to-sterilize locations (e.g. deep inside any crevice or packaging). Also place temperature/humidity sensors if needed. Timeline: month 6–7; Staff: QA Validation Eng, Manufacturing/Tech to load.
Microbiological PQ (EO Overkill method): Run three half-cycle EO runs . After each half exposure, remove BIs and send for incubation (typically 7 days) to check for any survivors. All BIs should show no growth to demonstrate the half-cycle achieved sterility assurance >10^-6 (overkill). In parallel, do one full-cycle run and actually test product sterility (may not be needed if overkill proven, but sometimes done as sterility test on product). Timeline: execution in month 7, BI incubation might run into month 8; Staff: Sterilization lab doing BI culture, RA monitors.
Physical PQ and Residuals: Run at least three full-cycle EO sterilization loads (with product) to ensure the process is consistent and the product isn’t damaged. After each, test product functionality (e.g. does it still fit/operate). Also, take samples for EO residual analysis – per ISO 10993-7, measure ethylene oxide, ethylene chlorohydrin, ethylene glycol residual levels in product to ensure they are below limits for patient contact devices. If above, might need additional aeration time or cycle adjustment. Timeline: month 7–8; Staff: Contract lab (chem analysis), RA review.
Reusable Device Reprocessing Validation: If applicable, soil devices with a standardized test soil, then have a cleaning procedure (manual or automated cleaning steps) performed, check cleanliness (protein residue etc.), then autoclave using hospital-type cycle. Do this for multiple cycles (e.g. 10 reprocessing cycles) to ensure device remains clean and functional and sterile. Also verify that no degradation (corrosion, etc.) after repeated autoclave. Timeline: month 7 (if done internally or at a reprocessing lab); Staff: QA, possibly outsource to reprocessing validation lab.
Sterilization Validation Report: Compile all results: include the cycle parameters, BI results (should be all sterile in half-cycle, which gives assurance), sterility test (if done), residual results, and conclusion that SAL 10^-6 is achieved. Also include if any product modifications were needed. If any failures occurred (like a BI growth), document the investigation and rerun after adjustments. For reusable, document that the cleaning/sterilization IFU was validated successfully (no visual soil, log reduction of microbes achieved, etc.). Timeline: month 8–9; Staff: QA Validation Engineer, RA.
Production Monitoring: After validation, implement routine monitoring for sterilization: e.g. batch records from contract sterilizer, periodic BI tests or dose audits (for gamma). Establish schedule for requalification (annual as per ISO 11135) . Timeline: post-launch (procedural control); Staff: QA.
Timeline & Resources: Sterilization validation must be completed before submission ideally, or at least in parallel (FDA expects sterilization info in 510(k)). So plan to finish by month 9. It’s often on critical path due to waiting times (BI incubation, etc.). Staffing: QA/Validation Engineer (0.3 FTE for 2–3 months) to coordinate and document, RA to interface with contract lab and ensure FDA documentation. Contract Sterilizer and Lab: significant external resources with fees. Tools: BI spore strips, data loggers, and laboratory incubators at the contract site. Possibly our own small autoclave for internal testing of reusable validation, plus protein test kits for cleaning efficacy.
Risk & Mitigation (Sterilization): Risks: Failure to achieve sterility in validation (BIs not all killed). Mitigation: Use overkill half-cycle approach with ample margin; if any BI survives, possibly product design is an issue (tight crevice) – could enlarge pathways or adjust cycle (e.g. more dwell, more humidity). Excessive EO residuals: Mitigate by adding enough aeration time in cycle (most EO providers do 12+ hours post-sterilization aeration), or use alternative sterilant if EO not feasible. Device damage from sterilization: e.g. plastic warping in autoclave or EO affecting material. Mitigate by material selection and doing a few trial sterilization cycles on prototypes early (we can do a couple hospital autoclaves cycles on a prototype to see effects). Schedule risk: EO validation can be time-consuming – mitigated by parallelizing where possible and close follow-up with vendor. The table below addresses sterilization validation risks:

## Packaging & Labeling Validation
Packaging & Labeling Validation | Unnamed: 1 | Unnamed: 2 | Unnamed: 3 | Unnamed: 4 | Unnamed: 5 | Unnamed: 6 | Unnamed: 7
Description & Standards: We need to ensure that the device’s packaging system reliably protects the product (and maintains sterility if applicable) through shipping, storage, and handling. This involves designing packaging (both the sterile barrier system for any sterile components and the transport packaging for the overall system) and then validating it per ISO 11607-1 & -2. ISO 11607-1 covers design of sterile packaging (material choice, shelf-life considerations) and ISO 11607-2 covers process validation for sealing and assembly of packaging . We will also perform distribution simulation testing (e.g. per ASTM D4169 or ISTA standards) – vibration, drop, compression – to ensure the shipping boxes protect the unit without damage. The packaging validation includes IQ/OQ/PQ of any sealing equipment (if we seal pouches or trays in-house) and integrity testing (dye leak, bubble leak tests) on sealed packages after environmental conditioning (temperature/humidity cycling) and simulated shipment. Labeling (content of labels and IFU) also needs review – we will ensure it meets FDA and ISO 15223/Unique Device ID requirements, and validate that labels remain affixed and legible after environmental tests. Key deliverables: Packaging Design Drawings/Specifications, Packaging Validation Plan, test reports for distribution, Seal Strength and Integrity Test Reports, Packaging Validation Report, and label proof approvals.
Packaging & Labeling: Risks, Harms & Mitigations
|  |  | Risk | Description | Potential Harm (if not addressed) | Mitigation | Key Deliverables
Steps & Tools: |  |  | Transit damage | Package fails (tears, crushes) under shipping stress (drop, vibration) – device or sterile barrier gets damaged | Device arriving non-functional or non-sterile at hospital, leading to delays in surgery or infection risk; customer complaints/returns | Use robust packaging design: ample cushioning foam around robot; design custom insert to immobilize parts; simulate worst-case (e.g. drop from taller height than spec as safety margin); conduct formative transit tests on mock-up packages and refine before official testing | Packaging Design Drawings (with cushioning details); Pre-test drop/vibe results and adjustments; Final ISTA Test Report (passing)
|  |  | Seal integrity failure | Sterile seal leaks (pinholes or channels) after packaging or post-transport | Sterility compromised – risk of infection; would cause field recall if found | Optimize sealing parameters via OQ (get strong, uniform seals); perform in-process seal inspections (seal prints, bubble tests on sample from each lot); design sealing jigs to ensure alignment; use adequate seal width; test seal integrity after environmental conditioning to verify robustness | OQ Report for Sealing (showing optimal settings); In-process QC plan (e.g. peel test each lot); Integrity test results (dye/bubble tests – all pass after conditioning)
Package Design & Materials: Design the primary packaging for sterile components (e.g. a double sterile barrier: product in a rigid tray with Tyvek lid or a sterilization pouch) that will maintain sterility. Also design secondary packaging: e.g. a kit box for instruments, and the shipping carton for the whole system (with cushioning, etc.). Select materials: Tyvek, medical-grade pouch, or tray that fits our parts; corrugated boxes, foam inserts for the robot base if shipping, etc. Ensure packaging materials are compatible with sterilization (Tyvek for EO, etc.). Timeline: month 5; Staff: Packaging Engineer or Mechanical Eng (packaging focus). |  |  | Aging effects on packaging | Over time (or with sterilization) packaging materials become brittle or labels illegible | Package could break in storage (loss of sterility) or user mis-uses if label unreadable (e.g. wrong expiration read) | Use materials with known stability (Tyvek has known 5-year stability); do accelerated aging and if any degradation seen, adjust shelf life accordingly; ensure label ink and adhesive are rated for the environment (UV, etc.); specify storage conditions if needed (temperature/humidity) in IFU | Accelerated Aging Test Report; Real-time aging plan (if needed for verification); Label adhesive spec sheet (with tested conditions); Labeled sample evaluation (no fading or detachment)
|  |  | Labeling content error | Mistake on label or IFU (wrong lot, missing info, wrong symbol, typos) | Regulatory non-compliance (potential 510(k) hold), or misuse by user (if instructions wrong) leading to patient harm | Implement thorough label review and approval process (RA, QA, engineering, and clinical review IFU to ensure clarity); use a checklist against 21 CFR 801 and ISO 15223 to ensure all required symbols/statements present; perform test printing to check UDI scanability and readability; consider user comprehension test for IFU critical steps | Approved Label Proofs (with review signatures); Labeling Checklist completed; IFU validation (usability feedback on instructions, possibly as part of summative usability to see if they can follow reprocessing instructions correctly)
Seal Process Development: If we will seal packages in-house (or at a vendor), identify sealing equipment (heat sealer for Tyvek pouches or tray sealer). Develop initial sealing parameters (temperature, pressure, dwell) to achieve good seals. Timeline: month 6; Staff: Manufacturing/Process Engineer. |  |  | UDI and regulatory marking | Failing to properly implement Unique Device Identifier or other required markings (CE mark, etc. if global) | For UDI: inability to track devices (recall issue); Regulatory issue (FDA can reject submission if UDI not in place, or later enforcement action) | Use labeling software to generate UDI barcode in correct format (GS1 or HIBCC); test scan it with a verifier for grade; ensure UDI data is submitted to GUDID database prior to marketing; keep abreast of global marking needs (though initial focus FDA); include required symbols (non-sterile if applicable, manufacturer address, etc.) per standards | UDI Label artwork; UDI Database submission confirmation (planned post-clearance); Symbol glossary in IFU (if needed for all symbols); Compliance statement in submission (e.g. “labels meet UDI requirements”)
|  |  | Packaging process validation | If sealing or assembly process not validated, there could be variability (some packages not sealed fully) | Some units in field could be non-sterile without detection, risking infection; also non-compliance with ISO 11607-2 | Complete IQ/OQ/PQ on sealing process: e.g. PQ by sealing multiple batches and testing seals from each to show consistency ; train personnel on sealing SOP; implement seal integrity checks each batch; monitor seal equipment calibration periodically | Sealer IQ/OQ Protocol & Report; Sealing Work Instruction; Training records for packaging operators; Packaging Process PQ report (consistent seal strength data)
Labeling Content Development: Draft all labels – device labels, sterile package labels, outer box labels – containing required info (UDI barcode, device name, lot/serial, expiration date if sterile, sterilization method, appropriate symbols per ISO 15223-1, and any handling icons like “Keep Dry”). Also prepare the Instructions for Use (IFU) document with necessary content (including the validated reprocessing instructions for reusables, and explanation of system usage, warnings). Circulate labeling for review and approval. Timeline: initial draft by month 6; Staff: RA (label content), Graphic Designer (layout).
Packaging Validation Plan: Write a protocol describing how we will validate packaging: include accelerated aging testing if shelf life needs to be claimed (e.g. we might want a 2-year sterile shelf life, so do accelerated aging at elevated temp for equivalent time), distribution simulation (vibration, drop, altitude, etc.), environmental conditioning (extremes of temperature/humidity), and package integrity tests (dye penetration or bubble leak on sterile seals). Also include seal strength tests (e.g. peel strength of seals via ASTM F88) to ensure adequate margin. Define acceptance criteria (no loss of integrity, no product damage). Timeline: month 6; Staff: QA/Packaging Engineer.
IQ/OQ of Packaging Equipment: Qualify any packaging machinery: e.g. heat sealer IQ (installed correctly, temp gauges calibrated) and OQ (establish sealing parameter range that yields acceptable seals – perhaps do a design of experiments varying temp/pressure to find optimal). Document this. If packaging is done by a contract packager, obtain their IQ/OQ documentation. Timeline: month 7; Staff: Process Validation Eng or Supplier.
Package Test Samples: Assemble test samples: sterilize them if applicable (to include effects of sterilization on packaging, e.g. EO could weaken seals slightly). Include worst-case – e.g. highest load in a box, or heaviest component at bottom for drop orientation. Condition these samples as per ISTA/ASTM guidelines (extreme hot/cold perhaps, and humidity). Timeline: month 7; Staff: Techs assembling test kits.
Distribution Simulation Testing: Send samples to a test lab or perform in-house if equipment exists, for vibration (shake on vibration table to simulate truck transport), shock (drop test from specified heights on all edges/corners – e.g. ISTA requires certain drops based on weight), compression (stacking load simulation), and if relevant, altitude/pressure (air transport simulation). After this, inspect packages and contents for damage. Timeline: month 7; Staff: Packaging Test Lab (external) or internal test engineer.
Seal Integrity & Strength Testing: On sterile packaged samples (post-distribution test and on some just aged), perform dye penetration tests (inject dye along seal to see if any leak path) or bubble leak test (submerge package under pressure to see bubbles) per ASTM F1608/F2096. Also do peel tests to measure seal strength and ensure it meets at least minimum (e.g. >1.0 N/15mm per internal spec) and consistency. Timeline: month 7–8; Staff: Packaging Lab or internal QA.
Accelerated Aging (if shelf-life claims): Place some packaged samples in an environmental chamber at elevated temperature (e.g. 55°C for ~~ to simulate 1–2 year real time). After aging, perform the same integrity tests. If real-time aging is needed to validate shelf life, we may do an initial accelerated to justify, and continue real-time in parallel (but submission can use accelerated results for interim). Timeline: start around month 6, readings by month 8; Staff: QA.
Label Adhesion/Legibility: As part of environmental conditioning, ensure labels remain stuck and readable (no fading or falling off) after humidity, etc. Do a rub test (per MIL-STD or ISTA) to verify print durability. Timeline: month 7–8; Staff: QA.
Packaging Validation Report: Compile all results: any damage from drops? (hopefully none, or acceptably minor scuffs), all seals intact (dye test pass), seal strength consistent, aging results acceptable. If a test failed (e.g. package burst on drop), address by redesign (add padding, stronger material) and repeat affected tests. Summarize that the packaging system meets ISO 11607 and that sterile barrier is maintained to the expiry. Timeline: month 8–9; Staff: QA/Packaging Engineer.
RACI for Packaging: (Just ensuring cross-functional involvement: e.g. RA approves labeling content, etc., which we handle in RACI matrix later.)
Timeline & Resources: Packaging development in month 5, testing in month 7–8, completion by month 9. It is somewhat on critical path due to aging time. Staffing: likely a Packaging Engineer or QA focusing on this (0.3–0.5 FTE). Use of an external package test lab for drops/vibration etc. Tools: Drop tester, vibration table, environmental chamber (often outsourced), dye leak test kit.
Risk & Mitigation (Packaging): Risks: Transit test failure – as Jeff Barrett noted, over a third of package validations fail transit tests (often due to not accounting for all shipping conditions). Mitigation: do feasibility tests early – e.g. drop a prototype package to see obvious issues; ensure proper cushioning and consider worst-case (like someone stacking it wrong). Seal integrity failure – mitigated by optimizing sealing process and verifying with tests on multiple samples; if failure, possibly redesign seal (wider seal flange, different material). Labeling errors – a very common cause of recalls (up to 40% of recalls from labeling ). Mitigation: rigorous label review process, and maybe get a label proof verification by scanning barcodes, checking content, and doing a mock regulatory review. UDI code must be correctly formatted – we’ll use a label software to generate and verify. Expiration date assignment – ensure it’s supported by aging data. Summaries in table:

## Usability & Human Factors Engin
Usability & Human Factors Engineering (IEC 62366 & FDA Guidance) | Unnamed: 1 | Unnamed: 2 | Unnamed: 3 | Unnamed: 4 | Unnamed: 5 | Unnamed: 6 | Unnamed: 7
Description & Standards: We will carry out a Usability (Human Factors) Engineering process to ensure the device is safe and effective when used by the intended users under expected conditions, per IEC 62366-1:2015 and FDA’s human factors guidance (Applying Human Factors and Usability Engineering to Medical Devices). This involves understanding users and use environments, identifying potential use errors and hazards, and iteratively improving the design to mitigate those. Ultimately, we will conduct a Summative Usability Validation (HF Validation Test) to confirm that the risk of use errors is acceptable (especially for tasks related to safety, often called Critical Tasks by FDA). We will maintain a Usability Engineering File documenting the process: User Profiles, Use Scenarios, Task Analysis, Hazard-Related Use Scenarios (per IEC 62366), and test reports. We will align with ISO 14971 by treating use errors that could lead to harm as hazards to be mitigated . Key deliverables: Use Specification, User Interface Requirements, Formative Study Reports, Summative (Validation) Test Protocol & Report, and traceability showing how use risks were mitigated.
Usability Engineering: Risks, Harms & Mitigations
|  |  | Risk | Description | Potential Harm (if not addressed) | Mitigation | Key Deliverables
Steps & Tools: |  |  | Critical use error not identified until late (or ever) | A use-related hazard is overlooked in analysis (e.g. user might mount a sensor incorrectly) and makes it to field | Could lead to serious incident (surgery error or injury) if not designed against; could necessitate field corrective action or addendum to IFU | Perform comprehensive task and use error analysis involving actual users or subject matter experts; include use scenarios in risk brainstorming; simulate out-of-normal scenarios (stress, time pressure) to see how user might err; cross-check with any known issues from similar devices (MAUDE database for use errors) | Task Analysis Document; Use Error Analysis (linked to risk file); updated risk controls addressing use errors (design or labeling changes)
|  |  | Inadequate formative testing | Design goes to final without sufficient user feedback – interface might be unintuitive | Users in OR may become confused or use device incorrectly, leading to delays or errors; adoption may suffer if device is hard to use | Conduct multiple formative tests on high-risk aspects (even informal ones with in-house staff or surrogate users if surgeons not available initially); use rapid prototypes (even paper or video prototypes for UI) to get early feedback; apply known human factors heuristics in design as baseline | Formative Test Reports (with findings and design changes tracked); Design review minutes showing HF input considered; Revised UI or hardware per feedback
Use Specification & User Profiles: Define the intended users (e.g. orthopedic surgeons, OR nurses/technicians) and their characteristics: skill level, training, physical attributes, etc. Also define use environment: an OR (bright lights, limited space, presence of sterile field, possibly time pressure). Document this in a Use Specification . Timeline: month 2; Staff: Human Factors (HF) Engineer or RA with clinical input. |  |  | Summative validation failure | Users in validation study commit serious errors or cannot complete tasks (failure criteria met) | Cannot confidently market device – likely regulators will require redesign/mitigation; project delay to fix issues and retest; if ignored and launched, high risk of use-error incidents | Mitigate by analyzing results in real-time during study – if an observation is trending poorly, consider pause and fix (if minor like changing prompt wording) before continuing; build time in schedule for potential re-testing; ensure training materials are adequate (though in validation, minimal training is given to truly test design) | Validation Test Report (if initial fails: record of issues and corrective actions; second test report showing success); Updated IFU or design modifications; possibly documented risk-benefit rationale if minor usability issues are accepted with training
|  |  | User training dependence | Device requires extensive training such that without it, use errors are likely (i.e. steep learning curve) | High risk if device is used by untrained personnel (maybe in emergency or turnover); could also limit market adoption because it’s too hard to learn | Simplify and automate where possible (design goal: make device as intuitive as possible); incorporate on-device guidance (contextual help, tutorials); plan comprehensive training program for all installs; in risk management, treat “use by untrained user” as misuse and provide mitigations (like labeling “For use by trained surgeons only”) | Training Materials; On-device prompts or quick reference guide; Labeling stating training requirements; Risk assessment on misuse by untrained addressed by distribution controls (only sell to trained)
Task Analysis: List out all tasks users must perform with the system, from setup, calibration, throughout the procedure, to cleanup. Identify critical tasks – those which, if done incorrectly or not done, could cause harm (e.g. entering patient data incorrectly, misaligning a reference, failing to clamp device resulting in drift). Also consider maintenance and reprocessing tasks for staff. Timeline: month 3; Staff: HF Engineer with design team. |  |  | IFU/Reprocessing instructions usability | Reprocessing steps (cleaning, etc.) or other IFU instructions not clear, leading to non-compliance by staff | E.g. improper cleaning could cause infection risk; misuse of device due to misunderstood instructions | Conduct usability testing on the IFU itself – have a few people follow IFU to perform a task (like assemble or reprocess) and see if they can do it correctly; include diagrams in IFU for clarity; keep language at appropriate reading level; have IFU professionally translated (if needed) to avoid ambiguity | IFU Validation results (part of summative or separate formative on IFU); Revised IFU with user-friendly format (sections, visuals); IFU meets readability standards (like 8th grade reading level if targeting broad users)
|  |  | Use environment factors | OR environment constraints (lighting, noise, workflow) not accounted for (e.g. screen not visible under OR lights, device footprint too large) | Device might not be usable in practice or causes OR workflow disruptions (leading to it being unused or causing delays) | From early design, consider OR setting: use high-contrast UI visible in bright light, ensure physical unit fits typical OR without blocking anything; take device to actual OR during formative evaluation if possible; consult OR nurses on setup practicality; test worst-case (small OR, crowded) scenarios | Environmental Requirements in specs (e.g. display brightness, device footprint and weight limits); OR Layout compatibility review; Summative test done in realistic OR setup confirming environment issues addressed
Hazard-Related Use Scenarios: Using the risk analysis and task list, determine which tasks have potential use errors that could lead to hazardous situations . For each, identify the potential use error (what user might do wrong, e.g. skip a step, misinterpret a prompt), possible reasons (slips, lapses, confusion, etc.), and the potential harm (these become hazard-related use scenarios per 62366). Record these in a use FMEA or use error analysis table. Timeline: month 3–4; Staff: HF Engineer, QA (tie into risk file).
User Interface Requirements & Prototyping: Based on identified use risks and user needs, specify requirements for the user interface (UI) and workflow (e.g. system must provide clear confirmation of registration success, must alert if tracking is lost, interface must be usable with gloves, etc.). Develop UI prototypes early – could be screen mockups, or even a UI simulation – and physical mockups for hardware (like control buttons placement). Tools: design software for UI, 3D prints for handles etc. Timeline: start in month 4; Staff: UX Designer, HF Engineer, Software UI developer.
Formative Evaluations: Conduct one or more formative usability tests with representative users (surgeons, etc.) using prototypes. For example, simulate a key portion of workflow (like setting up and registering the system on a dummy patient). Observe and interview users to identify any confusion, errors, or preferences. If possible, do this in a realistic setting (mock OR). Ensure to cover critical tasks in these tests. Use results to refine design – e.g. improve instructions, change a display, add a physical guard to prevent misuse. Document findings and changes made. Timeline: one round around month 5, possibly another by month 6 after design tweaks; Staff: HF Engineer (moderator), 3–5 users per round, design team to implement changes.
Usability Risk Mitigations: Incorporate mitigations for use errors: could be design changes (make a control intuitive, add an interlock so sequence can’t be done wrong), protective measures (warning on screen or physical labels), or user training/IFU improvements. Update the risk management and usability files to show how each hazardous use scenario is addressed. Timeline: ongoing by design freeze; Staff: HF + design team, QA.
Summative Usability (Validation) Test Planning: Develop a Human Factors Validation Test Protocol for the final design. Define test objectives: to validate that intended users can perform critical tasks safely and effectively without serious use errors or problems. Scenarios should represent actual clinical use (e.g. a simulated surgery using anatomic model or cadaver, where surgeon uses the nav system as they would normally). Define success criteria (e.g. 90% of users complete each critical task without assistance or error; any errors that do occur are not safety-critical or are recovered by user). Plan to have at least 15 participants (FDA typically expects ~15 per user group for validation) representative of intended users, in an environment mimicking OR. Timeline: protocol by month 7; Staff: HF Engineer, RA review.
IRB/Ethics (if needed): If using human subjects (even just clinicians in a simulated task), determine if IRB is needed – usually not for usability testing with device prototypes, but ensure participants sign consent and data is kept confidential. If cadaver labs used, follow those site rules.
Summative Usability Test Execution: Conduct the test (month 8 perhaps). Have users attempt to perform a full procedure or relevant portions with minimal training (use only the IFU or quick guide we provide). Observe and record any use errors, difficulties, or help requests. Particularly note if any critical task failures occur. Have an unbiased moderator (HF specialist) conduct and record. Use video if allowed for analysis. After tasks, administer a questionnaire (e.g. SUS for usability or device-specific survey) and brief interview to get subjective feedback. Ensure any use errors are categorized by severity (was there potential harm? did user recover?). Staff: HF Engineer (moderator), Clinical proctor for safety, 15 users (surgeons or relevant staff) across possibly 2–3 sessions.
Usability Validation Results & Mitigation: Analyze the summative test data. If all critical tasks were completed successfully by nearly all users, and any errors were minor or mitigated, we can conclude usability is acceptable. If not – e.g. multiple surgeons made a dangerous error (like misinterpreted a display leading to a localization error) – we must address it (either redesign and possibly re-test, or justify via risk-benefit if truly minimal risk). The goal is no unacceptable use-related risks remain . Prepare a Usability Validation Report detailing the test and results, including any residual issues. Timeline: analysis and report by month 9; Staff: HF Engineer, RA.
Integration into Submission: Place the Usability Engineering Report in the DHF and include a summary in the 510(k) (FDA often expects a brief summary of human factors if it’s critical to safety, especially if labeling says “user must be trained” etc.). Also ensure IFU includes any necessary user warnings or training info gleaned from this process.
Training Program (if needed): If the device requires user training to operate safely (likely surgeons will get training from our company), develop a training plan and materials (could be an on-site demo, user manual, videos). While not a formal part of design controls, it mitigates use risk and we might mention in risk files that adequate training is provided. Timeline: by launch; Staff: Clinical Educator, Marketing.
Timeline & Resources: Human factors runs throughout design. Formative testing in mid development (month 5–6), summative near the end (month 8). Must be completed before submission if any critical tasks related to safety (FDA wants results to review if they deem necessary). Staffing: 1 Human Factors Engineer/Specialist (either in-house or consultant) ~0.2–0.3 FTE across project. Need access to representative end-users (surgeons) – will likely incentivize via honorarium or coordinate with a hospital partner for usability studies. Tools: Possibly a simulated OR setting, video recording equipment, data capture forms.
Risk & Mitigation (Usability): Risks: Use error leading to harm in field because it wasn’t caught or mitigated. Mitigation: thorough HF process – iterative tests to catch those errors and design them out or provide safeties. Failing summative test (users make serious errors) – mitigated by doing enough formative testing and not rushing to validation until design is sound; if it fails, allocate time to fix and repeat test if needed (which could delay submission). User variability – mitigated by recruiting a diverse set of test participants (different experience levels, etc.). Also risk that usability file is incomplete – mitigated by following IEC 62366 steps and documenting all, and possibly having an HF expert peer review. Table below outlines key usability risks:

## Cybersecurity & Data Privacy
Cybersecurity & Data Privacy | Unnamed: 1 | Unnamed: 2 | Unnamed: 3 | Unnamed: 4 | Unnamed: 5 | Unnamed: 6 | Unnamed: 7
Description & Standards: Given the device likely involves software, possibly wireless connectivity (maybe to pull imaging data or send data to hospital network), and a computer controller, cybersecurity is a crucial aspect. The FDA now requires robust cybersecurity design and documentation for connected devices (per the 2023 FDA guidance “Cybersecurity in Medical Devices: Quality System Considerations and Content of Premarket Submissions”). We will incorporate a Secure Product Development Framework: conducting threat modeling, implementing cybersecurity controls (user authentication, encryption of data, secure boot, etc.), and planning for ongoing monitoring and updates. We will produce a Cybersecurity Risk Management Plan and include cybersecurity in the hazard analysis. We will also prepare documentation for submission: a Threat Assessment, Defensive Controls Implementation summary, Security Testing Report (e.g. results of vulnerability scans, penetration testing), and a Software Bill of Materials (SBOM) listing all third-party software components (OS, libraries) with versions . The SBOM helps in vulnerability management (identifying components with known vulnerabilities via CVEs). We will follow standards like FDA’s 2023 Guidance, NIST Cybersecurity Framework, and FDA’s 10 premarket cybersecurity principles. We must achieve compliance with new requirements under the FD&C Act Section 524B: including SBOM and plan to monitor and address vulnerabilities postmarket .
Cybersecurity: Risks, Harms & Mitigations
|  |  | Risk | Description | Potential Harm (if not addressed) | Mitigation | Key Deliverables
Steps & Tools: |  |  | Vulnerability leading to device hack | A weakness (open port, weak credential, buffer overflow) allows an attacker to alter device behavior or render it unusable (e.g. ransomware on device) | Could result in incorrect navigation output (patient injury), device unavailability (surgery delay), privacy breach of patient data | Secure development lifecycle: threat model, code reviews focusing on security; remove default passwords, enforce strong auth; use intrusion detection if feasible; conduct penetration testing to discover and fix vulnerabilities ; maintain SBOM to quickly patch known vulns ; apply FDA cybersecurity guidance fully | Threat Model Document; List of Cyber Controls implemented; Penetration Test Report (with all critical issues resolved); SBOM (for transparency and maintenance)
|  |  | Non-compliance with new FDA cyber requirements | Missing SBOM or inadequate documentation in 510(k) submission | FDA may refuse submission (RTA) or issue deficiency, delaying clearance significantly | Use FDA guidance checklist as template; ensure SBOM is complete and in recommended format; include a plan for updates and vulnerability handling as required by Section 524B ; possibly get a cybersecurity expert to review submission package before sending | Cybersecurity Documentation package (in submission, addressing all required elements); Internal review sign-off from security expert or consultant; SBOM included in submission and provided to users (if required)
Threat Modeling: Systematically analyze the system architecture to identify potential cyber threats. Use methods like STRIDE to consider Spoofing, Tampering, Repudiation, Information disclosure, Denial of service, Elevation of privilege. Involve software, IT, and clinical representatives. Identify assets (patient data, device control), entry points (USB ports, network connections), threat agents (malicious attacker, malware). Document threat scenarios (e.g. attacker could connect to network port and send false navigation data). Timeline: month 4; Staff: Software Architect or Security Engineer, Systems Engineer, QA. |  |  | Post-market vulnerability emergence | After launch, a new CVE is found in a third-party component used by the device (or new attack method) and device is exposed | If not patched, could be exploited in field leading to harm; also regulatory enforcement if we don’t respond (per new regulations, failure to remediate could result in penalties) | Have a post-market surveillance plan for cybersecurity: subscribe to CVE databases for our SBOM components; maintain a response team; follow Coordinated Vulnerability Disclosure practices (work with ISAOs perhaps); define patch release process (e.g. firmware updates delivered within X months of vulnerability discovery) ; inform users of mitigations interim if needed | Post-Market Cybersecurity Plan (part of QMS); Service bulletins or patch notes process; Customer communication template for vulnerabilities; Inclusion of this plan in submission as commitment
|  |  | Data privacy breach | Patient data (imaging, patient ID) stored on device could be accessed or exfiltrated by unauthorized party | HIPAA violation, patient privacy compromised, potential legal liability, loss of trust in device | Limit storage of sensitive data on device; encrypt any stored or transmitted PHI; require user authentication to view patient data; provide ability to wipe data; include privacy risk assessment; if cloud connectivity, ensure secure protocols and agreements in place | Data Flow Diagram with privacy considerations; Privacy Impact Assessment; Device settings for data deletion; Statement of HIPAA compliance approach (for US marketing)
Security Requirements: Derive and document cybersecurity requirements from threat model: e.g., “System shall authenticate users before enabling calibration,” “All network communication shall be encrypted (TLS 1.3)”, “Software updates shall be digitally signed,” etc. Also include requirements for physical security (locks, etc., if needed in hospital). These become part of design inputs. Timeline: month 4–5; Staff: Security Engineer, Software lead. |  |  | System availability (Denial of Service) | Device could be susceptible to DoS (e.g. network flood causes lag or crash) | Unavailable device during surgery -> could force revert to manual or cause surgical delay, risking patient safety if mid-procedure | Include network hardening: rate-limiting, input validation to avoid crashes; test with fuzzing and high traffic to ensure stability; if critical, device should have offline mode (e.g. not fully dependent on network); incorporate in risk controls that any failure leads to safe state (surgeon alerted and can continue manually) | DoS Test results (system sustains or gracefully handles); Risk mitigations documented (e.g. alarm on comms loss); Possibly an internal test report on high-load network conditions
|  |  | Lack of user cybersecurity guidance | Users (hospitals) not given guidance to secure the device environment (e.g. changing default password, network config) | Device might be installed with weak security (user doesn’t change default password, etc.), increasing risk of breach | In IFU or separate Cybersecurity User Guide, include best practices: password change procedure, recommended network settings (firewall ports, VLAN), how to apply updates; provide a cybersecurity checklist to hospital IT during installation; possibly implement forced password change on first use | Cybersecurity User Instructions (part of IFU or install guide); Installation guide for IT; Default credentials management (device forces change)
Implement Security Controls: During design and coding, implement the required controls:
User Authentication: e.g. require login for critical functions, or at least role-based access (maybe an authorized operator vs. limited mode).
Access Control: Limit access to debug ports, secure boot to prevent unauthorized firmware.
Data Encryption: Encrypt patient data stored or transmitted (HIPAA considerations if any PHI stored).
Network Hardening: Use only necessary network services, close unused ports, include firewall rules.
Software Hardening: use stack protections, memory protections in code. Include anti-malware if Windows-based.
Physical Protections: Tamper-evident seals on panels if needed.
Document each control and link to threats mitigated. Timeline: month 5–7; Staff: Software/IT team, Security consultant oversight.
Cybersecurity Testing: Plan and conduct tests focused on security:
Vulnerability scanning: Run static and dynamic analysis tools on the software (for known vulnerability patterns, open ports, etc.).
Penetration Testing: Hire an ethical hacker or use an internal team to simulate attacks on the device – attempt to break in, escalate privileges, inject false data. Particularly test any network interfaces (wired or wireless).
Fuzz Testing: Feed random or boundary data to communication interfaces to see if system crashes or becomes unstable, which could be exploited.
Account/Password policy testing: Try default credentials, weak passwords to ensure system enforces strong ones.
Ensure testing environment mimics the deployed setup (connect device to network, etc.). Timeline: month 7; Staff: External cybersecurity firm or internal security engineer.
Results & Hardening: Address any vulnerabilities found in testing: e.g. patch third-party libraries that scanner flagged (update to newer version), close any open ports that were unnecessary, improve input validation if fuzzing found a crash, etc. Iterate until tests show no critical vulnerabilities remain. Document unresolved minor vulnerabilities with rationale if any (maybe some low risk ones). Timeline: month 7–8; Staff: Software team.
SBOM Creation: Compile a comprehensive Software Bill of Materials – list all software components (OS, third-party libraries, open-source packages) with name, version, and any known vulnerabilities (CVE) status. Use tools or manual tracking. The SBOM is required in the submission and we will provide it in a machine-readable format if possible. We will also implement processes to keep SBOM updated with future patches (postmarket). Timeline: month 8; Staff: Software Config Manager.
Cybersecurity Documentation: Prepare the Premarket Cybersecurity Report per FDA guidance. This will include:
Device cybersecurity overview: device features, connections, data types.
Threat model summary and risk assessment: including risk ratings of each threat, and how we addressed them (controls or justification).
Summary of security controls: grouped by FDA’s recommended categories (identify, protect, detect, respond, recover).
Traceability matrix from each identified risk to control implemented.
Testing summary: results of pen test, etc.
SBOM as attachment.
Secure update plan: how will updates be provided (e.g. via service laptop or network update, with what authentication).
Postmarket plan: how we will monitor (subscribe to vulnerability alerts for our components, have a coordinated vulnerability disclosure program, patch timeline commitments).
Use FDA’s guidance checklist to ensure all points addressed (e.g. cyber PHA, device cybersecurity bill of materials, etc.) . Timeline: month 8–9; Staff: RA, Security Engineer.
Internal QMS updates: Ensure our QMS procedures include cybersecurity: e.g. design input includes security, risk management includes threats, and CAPA process covers security patches. Also, ensure supplier controls are in place for any software components (e.g. verify COTS OS is configured securely). Timeline: ongoing; Staff: QA.
Training & Access Controls: Internally, limit access to source code and signing keys (to avoid insider threats), train development team on secure coding. Document these QMS measures if needed.
Privacy Considerations: If the system handles patient data (imaging, registration data), ensure compliance with privacy laws (HIPAA if applicable). That means data encryption and possibly no retention of patient identifiers on device unless needed. Document data flows in our cybersecurity documentation.
Timeline & Resources: Start threat modeling by month 4, implement controls by design freeze, perform pen test by month 7–8, finalize documentation by month 9. Staffing: Possibly engage a Cybersecurity Consultant or hire (0.2 FTE) to lead threat modeling and pen testing, plus involvement of the Software team to implement controls (some additional effort but within their tasks) and RA to compile documentation. Tools: Threat modeling tools (Microsoft Threat Model), vulnerability scanners (Nessus, NMAP, etc.), static analysis tools.
Risk & Mitigation (Cybersecurity): Ransomware or remote attack risk – if the device is networked, an attacker could compromise it, leading to loss of function or bad guidance (patient harm). Mitigation: robust security design as above, including ensuring device fails safe (if compromised, it should shut down rather than do wrong action). Regulatory risk – new FDA rules allow RTA (refuse-to-accept) if cybersecurity documentation is insufficient (post-March 2023). Mitigation: follow the guidance to the letter, include SBOM, etc., and possibly do a pre-submission question to FDA if unsure. Postmarket risk – vulnerabilities discovered after deployment. Mitigation: have a response plan (maybe sign up for ISO 27001 or follow TIR97 for legacy device cybersecurity), commit to patch timelines. Table summarizing:

## Regulatory Submission & Documen
Regulatory Submission & Documentation | Unnamed: 1 | Unnamed: 2 | Unnamed: 3 | Unnamed: 4 | Unnamed: 5 | Unnamed: 6 | Unnamed: 7
Description: Prepare comprehensive Regulatory Documentation to support a FDA 510(k) submission (and other regulatory needs). This includes compiling the Design History File (DHF) for internal compliance and creating the 510(k) submission package with all required sections: device description, intended use, predicate comparison, summaries of nonclinical testing (safety, EMC, performance, software, biocomp, etc.), labeling, and the 510(k) summaries and forms. We will ensure all prior deliverables (design controls, risk management, etc.) feed into this documentation. Also, concurrently, maintain our ISO 13485 Quality Management System documentation (Design and Development outputs, verification reports, etc.) which will form the Technical File if needed for CE marking later. The 510(k) focuses on demonstrating substantial equivalence to a predicate device – thus we need a Predicate Comparison Table showing how our device is similar/different in terms of intended use, tech characteristics, etc., and justify any differences do not affect safety/effectiveness. We will prepare a robust 510(k) Summary and Truthful and Accuracy Statement, and address any specific FDA guidance applicable (e.g. for orthopedic navigation systems, any special controls or guidance documents – possibly the FDA guidance on orthopedic computer-assisted surgical systems if exists). Key deliverables: 510(k) Submission Dossier (with sections such as: Administrative Info, Device Description, Predicate Comparison, Tables of Components, Software Section, EMC/Safety Testing, Risk Analysis summary, Biocomp, Sterilization info, Labeling, etc.), and an internal Regulatory Strategy Document.
Regulatory Documentation: Risks, Harms & Mitigations
|  |  | Risk | Description | Potential Impact (if not addressed) | Mitigation | Key Deliverables
Steps & Tools: |  |  | Incomplete submission | Omitting required info or sections (e.g. missing test report, SBOM, or inadequate device description) | FDA will refuse to accept or later issue deficiency, causing delay in clearance (project launch delay, increased cost) | Use FDA’s RTA checklist line-by-line prior to submission; engage a third-party regulatory expert to QC the package; ensure all testing outlined in plan has a summary in submission (if a test is still ongoing, consider waiting or explain clearly and provide timeline if acceptable) | Completed FDA eSTAR 510(k) with all attachments; Internal RTA checklist signed off; If available, documentation of a pre-submission and FDA feedback to guide content
|  |  | Predicate mismatch | Chosen predicate not sufficiently similar or differences not well justified | FDA may say device is NSE (not substantially equivalent), requiring de novo or additional data – major delay | Choose predicate wisely (prefer one with same intended use and similar tech); if differences (like AI), provide robust data showing equal or better performance; possibly include multiple predicates if needed to cover features (but ensure not creating “combo” beyond allowed); clearly state no new questions of safety/effectiveness arise from differences | Predicate Device Comparison Table; Rationale for equivalence document; Performance test reports demonstrating equivalence; 510(k) Summary highlighting similarities
Regulatory Strategy & Predicate Identification: Early on (already done or to do in parallel), identify a suitable Predicate Device (e.g. an existing FDA-cleared surgical navigation or robotic system) and gather its 510(k) summary and, if possible, labeling. Assess differences: if our device has new features (like AI) ensure they are still within a 510(k) route (likely yes for Class II). Plan if any Special Controls apply (some devices have special controls in the classification). Document this strategy in a brief memo. Timeline: month 1–2; Staff: RA lead. |  |  | Documentation inconsistency | Data in one section doesn’t match another (e.g. IFU says one thing, device description says another, or test report values conflict with specs) | FDA confusion, questions (slowing review); potential loss of credibility of submission | Conduct final editorial review cross-check: ensure device specs are consistently stated (in description, IFU, testing); have someone not involved in writing do a read-through for consistency; update all documents to latest device version (especially if late changes occurred); control versions in QMS to avoid mix-ups | Final Device Description (with version that matches tested device); IFU aligned with design (reviewed after design freeze); Consistency checklist (matching values like dimensions, accuracy claims across docs)
|  |  | Labeling issues (regulatory) | IFU or labels not meeting regs (e.g. missing required warning, or UDI format error) | Could lead to post-clearance issues (483 observation or need to correct labels; in worst case, clearance could be delayed if labeling insufficient for safe use) | Have RA and QA thoroughly review labeling against FDA device labeling guidance and relevant standards; especially verify that all risk mitigations from risk management are either built in design or addressed in labeling (any “information for safety” needed is present); ensure IFU includes all reprocessing instructions validated, etc.; submit labeling in draft form in 510(k) and be prepared to tweak if FDA suggests changes | Labeling Review Checklist (with ISO 15223 and 21 CFR Part 801 compliance); Copy of IFU and labels in submission; Labeling addressing all residual risk (cross-checked with risk report)
Pre-Submission (if needed): Consider an FDA Q-Submission meeting to get feedback on novel aspects (like AI validation approach or cybersecurity approach). If done, schedule it by mid-project so FDA feedback can be incorporated. Timeline: maybe month 4; Staff: RA (prepares Q-sub package). |  |  | Extended FDA review | Unanticipated questions (e.g. FDA asks for additional testing or clarification not readily available) | Could push clearance beyond expected date, affecting project timelines and market entry | Mitigate via pre-submission for high-risk questions (like novel AI – get agreement on validation approach beforehand); maintain a rapid response team post-submission (engineering/clinical on standby to draft replies and perhaps run minor additional tests if absolutely required during review); good documentation quality to reduce questions; possibly have backup data (like extra stress test results) ready if likely to be asked | Pre-Sub meeting minutes (with FDA feedback to incorporate); Team assignments for possible queries (who will address software, who addresses clinical, etc.); All test raw data archived in case deeper analysis needed quickly
|  |  | QMS audit readiness | While preparing submission, neglecting QMS could lead to issues (e.g. if FDA comes for inspection around clearance time) | Observations or non-compliance findings; potential hold on clearance if quality system is inadequate for manufacturing phase | Continue robust QMS implementation: by submission time, have key procedures in place (CAPA, complaint handling, etc.), and ensure DHF and DMR are organized; optionally perform a mock audit or gap assessment to ISO 13485 and 21 CFR 820; address any gaps before production; emphasize design controls and risk management traceability in DHF is complete | Quality Manual and Procedures updated; Internal audit reports; DHF index with all required documents present; Management review minutes discussing readiness for manufacturing
Compile Technical Documentation: As each test/report is completed (biocomp, EMC, etc.), gather and format them for submission. Write summary narratives for each:
Device Description: Provide detailed diagrams of system components, photos, and an explanation of how it works (without revealing proprietary algorithm detail beyond needed). Include a high-level block diagram, and physical characteristics (dimensions, weight).
Indications for Use: Ensure we have a precise Indications statement, and it matches predicate’s scope to maintain equivalence.
Predicate Comparison: Table comparing key features (principle of operation, anatomy, accuracy, sterility, power, etc.), noting any differences and why they don’t raise new questions. If differences exist (e.g. we use AI), explain how performance testing addressed that.
Nonclinical Test Summary: Summarize all verification and validation testing:
Mechanical accuracy test results, mechanical safety (e.g. we might include results showing system’s accuracy and reliability).
Electrical safety & EMC: state compliance with standards (with report references).
Software: include Software section (Level of Concern doc, summary of development and verification – likely follow content of FDA’s software checklist with references to SRS, test reports, anomaly list).
Risk management: include a summary or list of identified hazards and mitigations (maybe a table of top risks and controls).
Usability/Human Factors: if any use risks were significant, mention that HF validation was done and successful (FDA often expects it if labeling or risk analysis suggests need).
Biocompatibility: reference our ISO 10993 testing and state outcomes (all passed).
Sterilization: state method, SAL 10^-6 achieved, EO residuals within limits, shelf life validated (if applicable).
Packaging: mention ISTA tests passed, etc.
Any specific performance tests: e.g. accuracy on bench, any cadaver studies if performed, etc., include those results.
Labeling: Attach copies of labels, IFU, any marketing brochures if available (FDA requires all labeling that will go out with device).
510(k) Summary: Prepare this last, summarizing device, predicate, testing, indications, etc., for public view.
Declarations/Forms: Fill Form FDA 3674 (clinicaltrials.gov not applicable likely, but must fill), 3602 form for small business if using, etc. Also a Truthful and Accurate statement on company letterhead.
Timeline: month 8–9 for writing and assembly; Staff: RA lead writer, engineers provide input, QA review.
Review & Approval of Submission: Conduct an internal review of the 510(k) dossier to ensure consistency and no gaps. Use a checklist from guidance (FDA Refuse-to-Accept checklist) to ensure everything is included. Get management sign-off. Timeline: end of month 9; Staff: RA, QA, possibly regulatory consultant for QC.
Submission to FDA: Submit via FDA’s eSTAR template (nowadays FDA encourages electronic submissions via their template). Monitor for RTA decision (they have 15 days to check completeness). If RTA issues, quickly address them. Then interact during Substantive Review (answer any questions, provide any additional data needed). Expect clearance ~3–6 months after submission. Timeline: submission target end of month 9 (assuming we started timeline around month 0 = now). Staff: RA liaison during review.
Quality System Documentation: In parallel, ensure all design control documents are signed and archived in DHF. Also compile the Device Master Record (DMR) (which includes product drawings, specifications, manufacturing instructions, BOM, etc.) in preparation for manufacturing. And prepare a Device History Record (DHR) for the first production units (the records of making verification units, perhaps). These are needed for internal compliance (FDA QSR inspection readiness). Timeline: by design transfer (month 9–10); Staff: QA, Manufacturing.
Regulatory Approvals Beyond FDA: While not immediately asked, if global launch is planned, begin assembling a Technical File for CE marking under MDR, which will include much of the above plus Clinical Evaluation Report, etc., but that might be out of scope here. We note it as a future task post-510(k).
Timeline & Resources: The submission writing happens primarily in the last 1–2 months once most testing done, though many sections can be drafted earlier. Aim to submit at end of Q3 in timeline. Staffing: Regulatory Affairs Specialist/Manager – primary (0.5 FTE during preparation, maybe 1 FTE for final month). Contributors: engineering/test team to provide data and review. Possibly a regulatory consultant to do a mock review. Tools: FDA eSTAR tool for 510(k), document templates for consistency, possibly a submission publishing software (if needed for formatting).
Risk & Mitigation (Regulatory Submission): Risks: Missing information leading to RTA refusal – mitigated by using FDA’s refuse-to-accept checklist and possibly pre-submission engagement. Substantive questions from FDA – mitigated by thorough testing and clear justifications (especially for anything novel like AI). We have to be prepared to answer quickly (have experts ready, maybe pre-draft explanations for expected questions, like how AI performance correlates clinically). Timeline risk – submission could slip if test data isn’t ready – mitigated by tracking all deliverables and having contingency (maybe submit with certain tests in progress only if agreed in pre-sub, though usually everything should be done). Also ensure U.S. Agent and establishment registration plans are set if needed. Table for regulatory documentation:

## Manufacturing Transfer & Proces
Manufacturing Transfer & Process Validation | Unnamed: 1 | Unnamed: 2 | Unnamed: 3 | Unnamed: 4 | Unnamed: 5 | Unnamed: 6 | Unnamed: 7
Description: As design development concludes, we transition to manufacturing. This involves scaling from prototypes to production-quality units, and validating that manufacturing processes produce devices that meet specifications consistently. Under ISO 13485 and FDA QSR, we must ensure Design Transfer is effective (design specs correctly translated to production instructions) . Key tasks: establish or work with manufacturing facilities (in-house assembly or contract manufacturer), qualify all production equipment (Installation Qualification), develop assembly and test procedures (with operator training), and perform Process Validation (OQ/PQ) for any processes that cannot be fully verified by subsequent inspection/testing. Likely processes needing validation include: sterile packaging sealing process, any sterilization process (covered earlier), perhaps software loading process (ensure correct version loaded each time), and any complex assembly or calibration processes. If injection molding or other special manufacturing is used for parts, those molds need OQ/PQ to ensure parts meet dimensional and mechanical specs. Also, we will validate any testing equipment (like final functional test rigs) to ensure they accurately verify the product (this can be considered part of OQ/PQ of test process). We will create a Production Control Plan (incoming inspection, in-process checks, final testing), and ensure all required IQ/OQ/PQ protocols are executed. After validation, the process enters routine production and we monitor via quality metrics (yield, etc.). Key deliverables: Manufacturing Plan, Standard Operating Procedures (SOPs) and Work Instructions for production, IQ/OQ/PQ protocols & reports for critical processes/equipment, Training Records for personnel, and a ready Device Master Record.
Manufacturing & Process Validation: Risks, Harms & Mitigations
|  |  | Risk | Description | Potential Harm (if not addressed) | Mitigation | Key Deliverables
Steps & Tools: |  |  | Low production yield / rework | Some assembly steps result in high defect rates (e.g. alignment step fails often, requiring re-adjust or scrap) | Could lead to delays in fulfilling orders, increased cost, or worse, devices with subtle defects passing through if not caught, affecting reliability in field | Improve jigs/fixtures to make alignment foolproof; add in-process checks earlier to catch issues before full assembly; provide additional training or refine work instructions; possibly refine design to be more assembly-tolerant (design for manufacturability improvements); monitor yield closely in initial production and perform CAPA if yield below target | Yield Data from pilot builds; Updated assembly procedure or jigs; Process FMEA (identifying high-risk steps and actions taken); CAPA records if yield issues and their resolution
|  |  | Process validation failure | During OQ/PQ, some process output not meeting spec (e.g. packaging seal strength occasionally below limit, or calibration process variability too high) | If unaddressed, could mean shipped product not consistently meeting specifications (safety/quality risk); regulatory non-compliance if process not validated but used | Investigate cause: e.g. equipment variation or human factor; tighten control (maintenance/calibration of equipment, or narrower parameter ranges); increase sample size in validation to be sure; after fixes, repeat OQ/PQ runs until criteria met; worst-case, implement 100% inspection if process inherently variable (while working on long-term fix) | OQ/PQ Reports (initial and repeat with improvements); Deviation/CAPA if initial validation failed, including root cause analysis and corrective action (e.g. equipment serviced or operator retrained); Final validation report showing all criteria met
Manufacturing Plan: Develop a plan outlining whether production is in-house or outsourced, key manufacturing steps, needed equipment and facilities, and required personnel skills. For example, assembly might be done in-house in an ISO 7 cleanroom (if needed for sterile components packaging), or contract out PCB assembly and then do final integration in-house. Plan should also identify any regulatory manufacturing requirements (e.g. if sterilization contractor needs to be ISO 13485 certified, etc.). Timeline: month 6; Staff: Manufacturing/Operations Manager. |  |  | Supplier part issues | Critical component from supplier does not meet spec (dimension out of tolerance, material off-grade) or is delayed | Could cause assembly issues, device performance degradation or delay production; quality issues could slip through causing field failures | Qualify suppliers thoroughly (audit, sample inspection); set up Incoming QC: measure critical dimensions or test functionality of incoming lots (with sampling plan); have backup suppliers identified for key parts; maintain some inventory buffer for long lead items; work with supplier on improvement if trends show drift | Supplier Quality Agreements; Incoming Inspection SOP and records; Approved Vendor List with dual sources for key parts; Nonconformance reports for any incoming rejection and supplier corrective actions
|  |  | Inadequate training / human error | Production staff make mistakes (miswire cable, wrong screw torque) due to lack of skill or unclear instruction | Device failures in field (if error not caught) or scrap/rework; potential safety risk if an assembly error defeats a redundancy | Implement training certification: each operator must assemble X units error-free under supervision before independent work; use visual aids and mistake-proofing (poka-yoke fixtures, color-coded connectors to prevent miswiring); implement in-process quality checks (peer inspection or automated test after critical steps); cultivate quality culture with operators (empower them to stop line if something seems off) | Training Records (signed off competencies for each operator); Assembly checklist that operators follow and sign for critical steps; Poka-yoke design descriptions (e.g. unique connectors only fit correct port); In-process inspection records (torque verification logs, etc.)
Supplier & Component Management: Finalize selection of suppliers for key components (e.g. motors, sensors, electronics). Ensure quality agreements in place and incoming inspection criteria set (e.g. test reports from PCB supplier, COC for materials). If any custom part is outsourced, perform supplier audits or quality assessments. Timeline: month 6–7; Staff: Supply Chain Manager, QA. |  |  | Traceability and mix-ups | Lot mix-up or missing trace (e.g. sterile lots not linked to BI results, or a component lot that fails later can’t be traced to devices) | Hard to execute recalls or pinpoint affected units if an issue arises; regulatory risk for QMS violation; potentially affected devices remain in market unknowingly | Establish lot traceability in production records: each device DHR records lot # of critical components (sensors, PCBs, etc.), sterilization lot, etc.; use barcode scanning or electronic systems to reduce manual errors; conduct mock recall exercise to ensure traceability system works (can we find all devices built with X lot?); maintain organized DHR files for each serial # device | Device History Record (DHR) template capturing all key trace fields; Implemented system (could be software or logs) for traceability; Mock recall test report demonstrating successful retrieval of unit list by lot; Procedure for lot segregation if nonconforming component identified
|  |  | Scaling up issues | The jump from prototype to volume production reveals capacity problems (single station too slow, one machine becomes bottleneck) | Unable to meet customer demand on time; or rushing could compromise quality | Evaluate takt time and cycle time for each process during pilot; if something is too slow, consider parallel stations or automation; use lean manufacturing principles to optimize flow (minimize wait times); have contingency plans for overtime or additional shifts if demand spikes unexpectedly at launch (until permanent capacity added) | Production Line Capacity Analysis (with cycle times); Action plan for bottlenecks (e.g. purchase second seal machine by X date if orders > Y); Lean layout diagrams or time-motion study results; Ramp-up schedule aligning workforce increase with expected demand
Device Master Record (DMR) Creation: Compile all documentation that defines the device’s production: drawings, component specifications, assembly instructions, software version/build instructions, BOM, labeling specs, and packaging specs. This DMR is essentially the “recipe” for manufacturing. Approve and release these documents under document control. Timeline: month 7; Staff: QA Document Control, Engineering provide content.
Pilot Build (Engineering Builds): Conduct one or more pilot production runs with the finalized design using production-intent components and processes. For example, build 5–10 units with manufacturing technicians following draft assembly instructions. Use this to identify any issues (parts that are hard to assemble, jigs needed, etc.). During pilot, also collect measurements or test data to establish baseline yields and to finalize any in-process test limits. Timeline: month 8; Staff: Manufacturing Engineers, technicians, overseen by R&D to troubleshoot.
Process Flow & Work Instructions: Define each process step (assembly step sequence, soldering, gluing, calibration, etc.). Write Work Instructions (WI) or SOPs with clear steps, photos/diagrams, and safety precautions (especially if any ESD or handling of sensitive parts). Include quality checkpoints (e.g. torque verification on screws, functional test after assembly). Timeline: month 7–8; Staff: Manufacturing Engineer, reviewed by QA.
Tooling and Equipment IQ: Install and set up any custom equipment: assembly jigs, torque drivers, test fixtures, sealing machines, sterilization chambers (if in-house), etc. Perform Installation Qualification (IQ): verify equipment installed as per specs, utilities correct, environment (e.g. temperature/humidity control for any sensitive assembly) is appropriate. Document model/serial of equipment, calibration certificates, etc. Timeline: month 8; Staff: Process Engineer/Technician, QA.
Operational Qualification (OQ): Develop protocols to test that each critical process produces acceptable output across its operating range. For example:
For a heat sealing machine (packaging): OQ by running samples at low, nominal, high settings of temp/pressure to confirm all within range yield good seals .
For injection molding: mold at different pressure/temperature to ensure dimensional output remains in spec range.
For assembly processes that rely on personnel: maybe not parametric, but do OQ by having different operators assemble to ensure instruction clarity.
For software loading: test that the software installation process reliably yields correct checksums on device.
For any computer-based test stations: OQ by simulating known good and bad units to ensure test detects failures reliably.
Document acceptance criteria (e.g. seal strength measurements, dimensional checks). Execute OQ and adjust process parameters to optimal. Timeline: month 8–9; Staff: Validation Engineer, manufacturing team.
Performance Qualification (PQ): Demonstrate that the process, at nominal settings with normal operators, consistently produces product meeting all specs, in a simulated production environment. Typically, run 3 consecutive lots of a defined size (e.g. 3 lots of 10 units each) and test them thoroughly to show all meet requirements. This would include final device testing on each unit. For sterile device, PQ included in sterilization validation already. For packaging, likewise 3 runs done. For assembly, PQ shows assembly yield is high and any defects are within acceptable rate. Document any non-conformance and investigation during PQ. Timeline: month 9; Staff: Manufacturing/QA.
Final Device Testing (Verification in Production): Establish final test procedure for each unit (could be same as some verification tests but done on each production unit). Validate any test software/hardware (this can be part of OQ/PQ). Ensure test equipment is calibrated and have IQ/OQ records. This ensures each shipped device is verified (traceability in DHR). Timeline: by month 9; Staff: Test Engineer, QA.
Training of Production Staff: Train assembly technicians on the SOPs and have them perform at least one supervised build. Document training records (each person qualified for specific process tasks). Also train QA inspectors on what to check (incoming inspection procedures, etc.). Timeline: month 9; Staff: Manufacturing Trainer or Engineer, QA.
Process Validation Report & Control Plan: Summarize the results of IQ/OQ/PQ for each validated process in a report. Also create a Control Plan listing each manufacturing step and the controls in place (inspection, test, monitoring, etc.), and any statistical process control if applied. Determine if any ongoing monitoring of process parameters is needed (e.g. record seal temperature each batch, etc.). Timeline: month 9–10; Staff: QA/Validation Engineer.
Pre-production Audit: Have QA do a walkthrough or audit of the manufacturing readiness before actual production: check that all procedures are in place, calibrated tools up-to-date, component supply and lot tracking system in place, etc. Address any findings. Timeline: end of month 9; Staff: QA auditor.
Launch Readiness: Once validation done and regulatory clearance obtained, formally release the manufacturing process for production. Ensure all required regulatory registrations (FDA device listing, establishment registration) are completed prior to shipping product.
Timeline & Resources: Manufacturing setup mainly in second half of project (months 6–10). Ideally, finish PQ around the time clearance is received so that distribution can begin promptly. Staffing: Manufacturing/Industrial Engineer (1 FTE coordinating transfer), Technicians/Operators for pilot and validation builds (maybe 2–3 people by end), QA Engineer (0.5 FTE ensuring validations and quality control plans), Supply Chain (0.2 FTE managing suppliers). Possibly Contract Manufacturer liaison if outsourcing (would coordinate in parallel). Tools: assembly jigs (designed by Mechanical eng or Mfg eng), calibrated tools (torque drivers, etc.), maybe MES (Manufacturing Execution Software) or at least travelers for DHR documentation.
Risk & Mitigation (Manufacturing Transfer): Risks: Production yield issues – something that was fine in prototype may be hard to assemble consistently (maybe tight tolerance fit leading to many rejects). Mitigation: pilot builds to catch issues, adjust tolerances or assembly method, train staff thoroughly, possibly invest in better tooling or gauge. Process validation failure – e.g. first PQ run shows a few units out of spec. Mitigation: analyze root cause (maybe operator error or slightly drifting machine parameter), refine process or tighten control on parameter, redo PQ runs until consistent. Supplier issues – part not meeting spec or late delivery. Mitigation: second-source critical parts, incoming QC to catch issues early, maintain some safety stock. Traceability and mix-ups – risk in production that wrong software version or wrong part gets into device. Mitigation: strong configuration management (each unit serial number, software loaded logged, use barcode scanning to ensure correct parts), line clearance procedures. Also maintain ESD controls to not damage sensitive electronics during assembly (train and equip with ESD mats etc.). Table:

## Project Timeline & Milestones
Project Timeline & Milestones
Month 0–1: Initiation & Planning – Regulatory strategy set (predicate identified), project plan approved, risk management plan drafted, user needs gathered. Milestone: Kickoff complete, Design & Development Plan approved.
Month 1–3: Requirements & Early Design – System requirements finalized; architecture design complete; initial risk analysis done; use specification and hazard use scenarios identified. Milestone: Requirements Freeze (end of month 3); System Arch Complete; Preliminary Hazard Analysis complete.
Month 3–5: Detailed Design & Alpha Prototyping – Mechanical CAD finalized for alpha, electronics schematics and PCBs layout done, software architecture defined, initial AI model trained; formative usability test #1 conducted; initial biocomp plan set; begin building alpha prototypes. Milestone: Design Review (Critical Design Review) at month 5 to green-light verification build.
Month 5–7: Verification Build & Testing – Build Beta prototypes using final design; perform design verification tests (mechanical accuracy, software verification, electrical safety pre-testing etc.); conduct formative usability test #2 if needed on near-final design; perform biocompatibility lab tests; start packaging/aging tests; start process validations (OQ). Milestone: Verification Testing Complete (end of month 7); Design Validation can start.
Month 7–8: Validation & Submission Prep – Conduct summative usability validation (month 8); execute sterilization validation (half cycles in month 7, complete by month 8); finish packaging distribution test; finalize all reports (verification, validation, RM report, HF report, etc.); compile 510(k) submission documents. Milestone: Design Validation Complete (e.g. summative usability pass); All documentation ready for review.
Month 9: 510(k) Submission & Manufacturing Validation – Submit 510(k) early in month 9; concurrently, execute manufacturing process PQ (3 lots); train production staff; do final QC of DMR/DHF. Milestone: 510(k) Acceptance by FDA (mid-month 9 if RTA passes); Manufacturing Process Validated.
Month 10–12: FDA Review & Launch Prep – Respond to FDA inquiries (months 10–11 if any); finalize inventory build of launch units (built under validated process); conduct mock recall and quality audit for launch readiness; achieve FDA Clearance by around month 12 (could be sooner ~month 11 if no major questions). Milestone: FDA 510(k) Clearance; Management review to authorize distribution.
Post-Launch (Beyond initial 12 mo): Ongoing post-market surveillance, further production ramp, and preparations for regulatory audits or international approvals.
Each task above has resource assignments as detailed in previous sections. Critical path items to monitor are software development/testing and regulatory testing (like biocomp and sterilization) due to their fixed durations.

## RACI
RACI Matrix (Responsibilities by Function) | Unnamed: 1 | Unnamed: 2 | Unnamed: 3 | Unnamed: 4 | Unnamed: 5 | Unnamed: 6 | Unnamed: 7 | Unnamed: 8 | Unnamed: 9
The table below assigns Responsible (R), Accountable (A), Consulted (C), and Informed (I) roles for major tasks across the key functions: PM (Project Management), SysEng (Systems Engineering), MechEng (Mechanical Engineering), ElecEng (Electronics Engineering), SoftEng (Software/AI Engineering), QA (Quality Assurance), RA (Regulatory Affairs), Clinical/HFE (Clinical Affairs/Human Factors), and Mfg (Manufacturing/Operations). This ensures clarity in ownership and communication:
Deliverable / Task | PM | SysEng | MechEng | ElecEng | SoftEng (incl. AI) | QA | RA | Clinical/HFE | Mfg
Project Plan & Scheduling | R (creates, maintains) | C (input on technical timelines) | I | I | I | C (quality milestones) | I | I | I
Design & Development Plan (ISO 13485) | R (develops) | C | C | C | C | A (approves compliance) | C | I | I
User Needs & System Requirements | C (coordinates collection) | R (owns requirements) | C | C | C | C (traceability check) | I | C (clinical input) | I
System Architecture Design | I | R | C | C | C | C | I | I | I
Risk Management Process (ISO 14971) | C | R (technical risk owner) | C | C | C | A (ensures process) & R (facilitator) | I | C (use risks) | I
Design Reviews (cross-functional) | R (organizes) | R (presents technical content) | R (for mech aspects) | R (elec) | R (soft) | C (ensure design control) | C (regulatory impact) | C (clinical usability) | I
Mechanical Design & Drawings | I | C (systems req alignment) | R (design parts) | I | I | C (DFM feedback, drawing control) | I | I | C (for manufacturability)
Electronics Hardware Design | I | C | I | R (PCB design) | I | C (verification input) | I | I | C (DFM, test fixtures)
Software Development & Verification (IEC 62304) | I | C (requirements) | I | I | R (implements & tests) | C (quality process) | C (submission docs) | I | I
AI Model Development & Validation | I | C (integration req) | I | I | R (data & model) | C (bias/risk check) | C (GMLP compliance) | C (clinical dataset input) | I
Biocompatibility Evaluation (ISO 10993) | I | I | R (material info) | I | I | C (risk assessment) | R (plans testing, coordinates lab) | I | I
Sterilization Validation (EO & Reuse) | I | I | R (provides samples, defines worst-case) | I | I | R (validation execution, IQ/OQ/PQ) | C (submission of sterilization info) | I | C (packing/sterile handling)
Packaging Design & Validation (ISO 11607) | I | I | R (package mech design) | I | I | R (validation testing, integrity) | C (label content regs) | I | R (seal process, shipping tests)
Usability/Human Factors (IEC 62366) | I | I | I | I | I (UI implementation) | C (risk of use) | I (review IFU) | R (plan & execute HF studies) | I
Cybersecurity Risk Management & Testing | I | R (systems threats) | I | I | R (implement controls, fixes) | C (risk file update) | C (documentation) | I | I
510(k) Submission Preparation | I | I | I | I | I | C (provide DHF docs) | R (author submission) & A (sign-off) | C (clinical sections) | I
Design Transfer to Manufacturing | I | C (tech clarification) | C (handover drawings/BOM) | C (handover specs) | C (handover build of materials) | C (ensure DMR complete) | I | I | R (owns transfer) & A (production readiness)
Manufacturing Process Validation (IQ/OQ/PQ) | I | I | C (tool design input) | C (test equipment design) | C (software loading) | R (protocols & quality sign-off) | I | I | R (executes validation runs)
Production Quality Control Plan | I | I | I | I | I | R (defines inspections/tests) | C (traceability regs) | I | R (implements on line)
Clinical Evaluation (if any, outside 510k) | I | I | I | I | I | I | R (if needed for CE or future) | R (for any clinical studies) | I
Post-Market Surveillance & Updates | I | I | I | I | I | R (monitor quality data) | R (reporting to regulators) | C (user feedback) | I (for process improvements)
In the RACI chart above, R = Responsible (executes the task), A = Accountable (ultimate decision-maker/approver), C = Consulted (provides input), I = Informed (kept updated on progress). For example, QA is Accountable for the overall risk management process being compliant, while Systems/Engineering is Responsible for identifying and mitigating technical risks; RA is Responsible for writing the 510(k) and Accountable for its completeness, but will Consult engineering for technical content. This matrix ensures every critical activity has clear ownership and cross-functional involvement, supporting a successful and compliant development process.

## Sheet18
Sources:
ISO 13485:2016 design and development planning requirements
FDA 21 CFR 820.30 design control regulations
Arterex Medical on design control phases and DHF
ISO 14971:2019 risk management traceability expectations
Greenlight Guru on risk controls linking to design inputs & verification
Emergo on biocompatibility within risk management process
Nelson Labs on ISO 10993-12 sample testing of final devices
Greenlight Guru on sterilization validation IQ/OQ/PQ and half-cycle method
Greenlight Guru podcast notes on packaging validation best practices (ISO 11607)
Medical Device HQ on usability engineering process per IEC 62366
FDA/HC/MHRA AI Good Machine Learning Practice principles
UL paper on IEC 80601-2-77 for surgical robots
Atlant Security on FDA cybersecurity requirements (risk plan & SBOM)
Vantage MedTech on FDA 2023 cyber guidance scope

## Project Plan
ID | Task Name | Area (Tab) | Responsible Party | Duration (weeks) | Duration (days) | Depends On (IDs) | Start Date | Finish Date | Leveled Start | Leveled Finish | Resource Delay (days)
1 | Draft Design & Development Plan | Design Controls & Documentation |  | 1 | 7 |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
2 | Define User Needs & System Requirements | Design Controls & Documentation |  | 4 | 28 | 1 | 2025-10-27 | 2025-11-24 | 2025-10-27 | 2025-11-24 | 0
3 | Develop System Architecture & Specifications | Design Controls & Documentation |  | 3 | 21 | 2 | 2025-11-24 | 2025-12-15 | 2025-11-24 | 2025-12-15 | 0
4 | Plan Design Reviews (PDR, CDR) | Design Controls & Documentation |  | 1 | 7 | 3 | 2025-12-15 | 2025-12-22 | 2025-12-15 | 2025-12-22 | 0
5 | Create Verification Plan (trace to requirements) | Design Controls & Documentation |  | 2 | 14 | 3 | 2025-12-15 | 2025-12-29 | 2025-12-15 | 2025-12-29 | 0
6 | Generate Design Outputs (CAD, Schematics, Code baselines) | Design Controls & Documentation |  | 8 | 56 | 2025-05-03 | 2025-10-20 | 2025-12-15 | 2025-10-20 | 2025-12-15 | 0
7 | Execute Design Verification (bench/system) | Design Controls & Documentation |  | 6 | 42 | 5 | 2025-12-29 | 2026-02-09 | 2025-12-29 | 2026-02-09 | 0
8 | Design Validation (simulated use) | Design Controls & Documentation |  | 4 | 28 | 7 | 2026-02-09 | 2026-03-09 | 2026-02-09 | 2026-03-09 | 0
9 | Design Transfer & DMR Compilation | Design Controls & Documentation |  | 4 | 28 | 8 | 2026-03-09 | 2026-04-06 | 2026-03-09 | 2026-04-06 | 0
10 | Create Risk Management Plan (ISO 14971) | Risk Management |  | 1 | 7 |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
11 | Preliminary Hazard Analysis (PHA) | Risk Management |  | 2 | 14 | 10 | 2025-10-27 | 2025-11-10 | 2025-10-27 | 2025-11-10 | 0
12 | System/Design FMEAs | Risk Management |  | 4 | 28 | 11 | 2025-11-10 | 2025-12-08 | 2025-11-10 | 2025-12-08 | 0
13 | Define Risk Controls & update requirements | Risk Management |  | 2 | 14 | 12 | 2025-12-08 | 2025-12-22 | 2025-12-08 | 2025-12-22 | 0
14 | Verify Risk Controls Implemented (tests) | Risk Management |  | 2 | 14 | 13 | 2025-12-22 | 2026-01-05 | 2025-12-22 | 2026-01-05 | 0
15 | Residual Risk Evaluation & Risk Report | Risk Management |  | 1 | 7 | 14 | 2026-01-05 | 2026-01-12 | 2026-01-05 | 2026-01-12 | 0
16 | Software Development Plan (IEC 62304) | Software Development |  | 1 | 7 |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
17 | Software Requirements Specification (SRS) | Software Development |  | 3 | 21 | 16 | 2025-10-27 | 2025-11-17 | 2025-10-27 | 2025-11-17 | 0
18 | Software Architecture & Detailed Design | Software Development |  | 3 | 21 | 17 | 2025-11-17 | 2025-12-08 | 2025-11-17 | 2025-12-08 | 0
19 | Implementation (Coding) | Software Development |  | 8 | 56 | 18 | 2025-12-08 | 2026-02-02 | 2025-12-08 | 2026-02-02 | 0
20 | Unit Testing | Software Development |  | 2 | 14 | 19 | 2026-02-02 | 2026-02-16 | 2026-02-02 | 2026-02-16 | 0
21 | Integration Testing (HW/SW-in-the-loop) | Software Development |  | 3 | 21 | 20 | 2026-02-16 | 2026-03-09 | 2026-02-16 | 2026-03-09 | 0
22 | Software Verification (System Tests) | Software Development |  | 4 | 28 | 21 | 2026-03-09 | 2026-04-06 | 2026-03-09 | 2026-04-06 | 0
23 | Software Validation (simulated use) | Software Development |  | 3 | 21 | 22 | 2026-04-06 | 2026-04-27 | 2026-04-06 | 2026-04-27 | 0
24 | Release & Configuration Audit | Software Development |  | 1 | 7 | 23 | 2026-04-27 | 2026-05-04 | 2026-04-27 | 2026-05-04 | 0
25 | Data Collection & Curation | AI Model Development & Validati |  | 3 | 21 |  | 2025-10-20 | 2025-11-10 | 2025-10-20 | 2025-11-10 | 0
26 | Dataset Split & Management | AI Model Development & Validati |  | 1 | 7 | 25 | 2025-11-10 | 2025-11-17 | 2025-11-10 | 2025-11-17 | 0
27 | Model Development & Training | AI Model Development & Validati |  | 5 | 35 | 26 | 2025-11-17 | 2025-12-22 | 2025-11-17 | 2025-12-22 | 0
28 | Internal Verification (metrics, robustness, bias) | AI Model Development & Validati |  | 2 | 14 | 27 | 2025-12-22 | 2026-01-05 | 2025-12-22 | 2026-01-05 | 0
29 | System Integration of Model | AI Model Development & Validati |  | 2 | 14 |  | 2025-10-20 | 2025-11-03 | 2025-10-20 | 2025-11-03 | 0
30 | Clinical/Simulated Performance Validation | AI Model Development & Validati |  | 4 | 28 | 29 | 2025-11-03 | 2025-12-01 | 2025-11-03 | 2025-12-01 | 0
31 | Algorithm Description & Documentation | AI Model Development & Validati |  | 1 | 7 | 30 | 2025-12-01 | 2025-12-08 | 2025-12-01 | 2025-12-08 | 0
32 | Schematic Design & Component Selection | Electronics Hardware Developmen |  | 3 | 21 |  | 2025-10-20 | 2025-11-10 | 2025-10-20 | 2025-11-10 | 0
33 | PCB Layout | Electronics Hardware Developmen |  | 3 | 21 | 32 | 2025-11-10 | 2025-12-01 | 2025-11-10 | 2025-12-01 | 0
34 | Prototype PCB Fabrication & Bring-up | Electronics Hardware Developmen |  | 4 | 28 | 33 | 2025-12-01 | 2025-12-29 | 2025-12-01 | 2025-12-29 | 0
35 | Bench/Engineering Tests | Electronics Hardware Developmen |  | 2 | 14 | 34 | 2025-12-29 | 2026-01-12 | 2025-12-29 | 2026-01-12 | 0
36 | EMC Pre-compliance Testing | Electronics Hardware Developmen |  | 1 | 7 | 35 | 2026-01-12 | 2026-01-19 | 2026-01-12 | 2026-01-19 | 0
37 | 60601 Safety Design Review | Electronics Hardware Developmen |  | 1 | 7 | 36, 35 | 2026-01-19 | 2026-01-26 | 2026-01-19 | 2026-01-26 | 0
38 | Design Freeze (Electronics) | Electronics Hardware Developmen |  | 1 | 7 | 36, 37 | 2026-01-26 | 2026-02-02 | 2026-01-26 | 2026-02-02 | 0
39 | Formal 60601/EMC/80601 Testing (3rd party) | Electronics Hardware Developmen |  | 3 | 21 | 38 | 2026-02-02 | 2026-02-23 | 2026-02-02 | 2026-02-23 | 0
40 | Conceptual Mechanical Design | Mechanical Design & Engineering |  | 3 | 21 |  | 2025-10-20 | 2025-11-10 | 2025-10-20 | 2025-11-10 | 0
41 | 3D CAD Modeling & Assemblies | Mechanical Design & Engineering |  | 4 | 28 | 40 | 2025-11-10 | 2025-12-08 | 2025-11-10 | 2025-12-08 | 0
42 | FEA & Tolerance Analysis | Mechanical Design & Engineering |  | 2 | 14 | 41 | 2025-12-08 | 2025-12-22 | 2025-12-08 | 2025-12-22 | 0
43 | Prototype Fabrication (Alpha) | Mechanical Design & Engineering |  | 4 | 28 | 42, 41 | 2025-12-22 | 2026-01-19 | 2025-12-22 | 2026-01-19 | 0
44 | Mechanical Verification (accuracy, ROM, load) | Mechanical Design & Engineering |  | 3 | 21 | 43 | 2026-01-19 | 2026-02-09 | 2026-01-19 | 2026-02-09 | 0
45 | Design Iteration (Beta) | Mechanical Design & Engineering |  | 2 | 14 |  | 2025-10-20 | 2025-11-03 | 2025-10-20 | 2025-11-03 | 0
46 | Finalize Drawings & DMR | Mechanical Design & Engineering |  | 2 | 14 | 45 | 2025-11-03 | 2025-11-17 | 2025-11-03 | 2025-11-17 | 0
47 | Identify Contact Materials & Classification | Biocompatibiliity & Material Sa |  | 1 | 7 |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
48 | Biocompatibility Evaluation Plan (BEP) | Biocompatibiliity & Material Sa |  | 2 | 14 | 47 | 2025-10-27 | 2025-11-10 | 2025-10-27 | 2025-11-10 | 0
49 | Select GLP Test Lab & SOW | Biocompatibiliity & Material Sa |  | 1 | 7 | 48 | 2025-11-10 | 2025-11-17 | 2025-11-10 | 2025-11-17 | 0
50 | Prepare Final-finish Samples (incl. sterilization) | Biocompatibiliity & Material Sa |  | 2 | 14 | 49 | 2025-11-17 | 2025-12-01 | 2025-11-17 | 2025-12-01 | 0
51 | ISO 10993 Testing (cytotox, sensitization, irritation, etc.) | Biocompatibiliity & Material Sa |  | 8 | 56 | 50 | 2025-12-01 | 2026-01-26 | 2025-12-01 | 2026-01-26 | 0
52 | Biocompatibility Report | Biocompatibiliity & Material Sa |  | 2 | 14 |  | 2025-10-20 | 2025-11-03 | 2025-10-20 | 2025-11-03 | 0
53 | Select Sterilization Method & Worst-case Load | Sterilization Process Validatio |  | 1 | 7 |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
54 | Sterilization Validation Plan (IQ/OQ/PQ) | Sterilization Process Validatio |  | 2 | 14 | 53 | 2025-10-27 | 2025-11-10 | 2025-10-27 | 2025-11-10 | 0
55 | Contract Sterilizer Coordination (schedule) | Sterilization Process Validatio |  | 1 | 7 | 54 | 2025-11-10 | 2025-11-17 | 2025-11-10 | 2025-11-17 | 0
56 | IQ/OQ Documentation (from vendor/in-house) | Sterilization Process Validatio |  | 2 | 14 | 55 | 2025-11-17 | 2025-12-01 | 2025-11-17 | 2025-12-01 | 0
57 | PQ: Half-cycle BI Runs (3x) & Full-cycle | Sterilization Process Validatio |  | 4 | 28 | 56 | 2025-12-01 | 2025-12-29 | 2025-12-01 | 2025-12-29 | 0
58 | EO Residuals / Physical PQ & Functional Checks | Sterilization Process Validatio |  | 2 | 14 | 57 | 2025-12-29 | 2026-01-12 | 2025-12-29 | 2026-01-12 | 0
59 | Sterilization Validation Report | Sterilization Process Validatio |  | 1 | 7 | 58 | 2026-01-12 | 2026-01-19 | 2026-01-12 | 2026-01-19 | 0
60 | Package & Label Design | Packaging & Labeling Validation |  | 2 | 14 |  | 2025-10-20 | 2025-11-03 | 2025-10-20 | 2025-11-03 | 0
61 | Sealing Process Development | Packaging & Labeling Validation |  | 1 | 7 | 60 | 2025-11-03 | 2025-11-10 | 2025-11-03 | 2025-11-10 | 0
62 | Packaging Validation Protocol | Packaging & Labeling Validation |  | 1 | 7 | 61 | 2025-11-10 | 2025-11-17 | 2025-11-10 | 2025-11-17 | 0
63 | Build Test Samples (sterilized) | Packaging & Labeling Validation |  | 1 | 7 | 62 | 2025-11-17 | 2025-11-24 | 2025-11-17 | 2025-11-24 | 0
64 | ISTA Transport Simulation & Integrity Tests | Packaging & Labeling Validation |  | 3 | 21 | 63 | 2025-11-24 | 2025-12-15 | 2025-11-24 | 2025-12-15 | 0
65 | Accelerated Aging (if shelf-life claimed) | Packaging & Labeling Validation |  | 12 | 84 | 64, 63 | 2025-12-15 | 2026-03-09 | 2025-12-15 | 2026-03-09 | 0
66 | Label Legibility/Adhesion Testing | Packaging & Labeling Validation |  | 2 | 14 | 65, 60 | 2026-03-09 | 2026-03-23 | 2026-03-09 | 2026-03-23 | 0
67 | Packaging Validation Report | Packaging & Labeling Validation |  | 1 | 7 | 66, 64 | 2026-03-23 | 2026-03-30 | 2026-03-23 | 2026-03-30 | 0
68 | Use Specification & User Profiles | Usability & Human Factors Engin |  | 1 | 7 |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
69 | Task Analysis & Critical Tasks | Usability & Human Factors Engin |  | 1 | 7 | 68 | 2025-10-27 | 2025-11-03 | 2025-10-27 | 2025-11-03 | 0
70 | UI/Workflow Prototyping | Usability & Human Factors Engin |  | 2 | 14 | 69 | 2025-11-03 | 2025-11-17 | 2025-11-03 | 2025-11-17 | 0
71 | Formative Usability Studies | Usability & Human Factors Engin |  | 2 | 14 | 70 | 2025-11-17 | 2025-12-01 | 2025-11-17 | 2025-12-01 | 0
72 | Summative Usability Test Protocol | Usability & Human Factors Engin |  | 2 | 14 | 71 | 2025-12-01 | 2025-12-15 | 2025-12-01 | 2025-12-15 | 0
73 | Summative Usability Test Execution | Usability & Human Factors Engin |  | 2 | 14 | 72 | 2025-12-15 | 2025-12-29 | 2025-12-15 | 2025-12-29 | 0
74 | Usability Validation Report | Usability & Human Factors Engin |  | 1 | 7 | 73 | 2025-12-29 | 2026-01-05 | 2025-12-29 | 2026-01-05 | 0
75 | Threat Modeling & Cyber Risk Assessment | Cybersecurity & Data Privacy |  | 1 | 7 |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
76 | Security Requirements (auth, encryption, SBOM) | Cybersecurity & Data Privacy |  | 1 | 7 | 75 | 2025-10-27 | 2025-11-03 | 2025-10-27 | 2025-11-03 | 0
77 | Implement Security Controls | Cybersecurity & Data Privacy |  | 4 | 28 |  | 2025-10-20 | 2025-11-17 | 2025-10-20 | 2025-11-17 | 0
78 | Vulnerability Scans & Pen Testing | Cybersecurity & Data Privacy |  | 2 | 14 | 77 | 2025-11-17 | 2025-12-01 | 2025-11-17 | 2025-12-01 | 0
79 | Remediation & Hardening | Cybersecurity & Data Privacy |  | 2 | 14 | 78 | 2025-12-01 | 2025-12-15 | 2025-12-01 | 2025-12-15 | 0
80 | SBOM Creation | Cybersecurity & Data Privacy |  | 1 | 7 | 79 | 2025-12-15 | 2025-12-22 | 2025-12-15 | 2025-12-22 | 0
81 | Cybersecurity Documentation (FDA) | Cybersecurity & Data Privacy |  | 2 | 14 | 80 | 2025-12-22 | 2026-01-05 | 2025-12-22 | 2026-01-05 | 0
82 | Regulatory Strategy & Predicate Selection | Regulatory Submission & Documen |  | 1 | 7 |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
83 | 510(k) Content Outline & Gap List | Regulatory Submission & Documen |  | 1 | 7 | 82 | 2025-10-27 | 2025-11-03 | 2025-10-27 | 2025-11-03 | 0
84 | Draft Device Description & Indications for Use | Regulatory Submission & Documen |  | 1 | 7 | 83 | 2025-11-03 | 2025-11-10 | 2025-11-03 | 2025-11-10 | 0
85 | Compile Nonclinical Test Summaries | Regulatory Submission & Documen |  | 2 | 14 | 84 | 2025-11-10 | 2025-11-24 | 2025-11-10 | 2025-11-24 | 0
86 | Labeling/IFU Drafts | Regulatory Submission & Documen |  | 1 | 7 | 85, 84 | 2025-11-24 | 2025-12-01 | 2025-11-24 | 2025-12-01 | 0
87 | 510(k) Summary & eSTAR Assembly | Regulatory Submission & Documen |  | 2 | 14 | 85, 86 | 2025-12-01 | 2025-12-15 | 2025-12-01 | 2025-12-15 | 0
88 | Internal Submission Review & Approval | Regulatory Submission & Documen |  | 1 | 7 | 87 | 2025-12-15 | 2025-12-22 | 2025-12-15 | 2025-12-22 | 0
89 | Submit 510(k) | Regulatory Submission & Documen |  | 0.2 | 2 | 88 | 2025-12-22 | 2025-12-24 | 2025-12-22 | 2025-12-24 | 0
90 | Manufacturing Plan & DMR Completion | Manufacturing Transfer & Proces |  | 2 | 14 |  | 2025-10-20 | 2025-11-03 | 2025-10-20 | 2025-11-03 | 0
91 | Pilot Build (Engineering Lot) | Manufacturing Transfer & Proces |  | 4 | 28 | 90 | 2025-11-03 | 2025-12-01 | 2025-11-03 | 2025-12-01 | 0
92 | Define Process Flow & Work Instructions | Manufacturing Transfer & Proces |  | 2 | 14 | 91, 90 | 2025-12-01 | 2025-12-15 | 2025-12-01 | 2025-12-15 | 0
93 | Tooling/Equipment IQ | Manufacturing Transfer & Proces |  | 2 | 14 | 92 | 2025-12-15 | 2025-12-29 | 2025-12-15 | 2025-12-29 | 0
94 | Process OQ | Manufacturing Transfer & Proces |  | 2 | 14 | 93 | 2025-12-29 | 2026-01-12 | 2025-12-29 | 2026-01-12 | 0
95 | Process PQ (3 lots) | Manufacturing Transfer & Proces |  | 4 | 28 | 94 | 2026-01-12 | 2026-02-09 | 2026-01-12 | 2026-02-09 | 0
96 | Train Production Personnel | Manufacturing Transfer & Proces |  | 2 | 14 | 95, 92 | 2026-02-09 | 2026-02-23 | 2026-02-09 | 2026-02-23 | 0
97 | Process Validation Report | Manufacturing Transfer & Proces |  | 1 | 7 | 96, 95 | 2026-02-23 | 2026-03-02 | 2026-02-23 | 2026-03-02 | 0

## Project Plan (By Date)
Project Plan for Class II Surgical Robotic Navigation System (510(k)) | ID | Task Name | Area (Tab) | Duration (weeks) | Depends On | Depends On (IDs) | Responsible Party | Start Date | Finish Date | Duration (days) | Leveled Start | Leveled Finish | Resource Delay (days)
| 1 | Draft Design & Development Plan | Design Controls & Documentation | 1 |  |  |  | 2025-10-20 | 2025-10-27 | 7 | 2025-10-20 | 2025-10-27 | 0
| 6 | Generate Design Outputs (CAD, Schematics, Code baselines) | Design Controls & Documentation | 8 | Create Verification Plan (trace to requirements), Develop System Architecture & Specifications | 2025-05-03 |  | 2025-10-20 | 2025-12-15 | 56 | 2025-10-20 | 2025-12-15 | 0
| 10 | Create Risk Management Plan (ISO 14971) | Risk Management | 1 |  |  |  | 2025-10-20 | 2025-10-27 | 7 | 2025-10-20 | 2025-10-27 | 0
| 16 | Software Development Plan (IEC 62304) | Software Development | 1 |  |  |  | 2025-10-20 | 2025-10-27 | 7 | 2025-10-20 | 2025-10-27 | 0
Software Development: Tasks include Create Software Development Plan (≈1 week) and Develop Software Requirements Spec (SRS) (≈2 weeks). After requirements, Software Architecture & Design (≈2 weeks) is done. Then Implementation (Coding) (≈4–6 weeks) follows, depending on architecture completion. Next are Unit Testing (≈2 weeks) and Integration Testing (≈3 weeks). Finally Software Verification & Validation (≈3 weeks, including system and regression tests). Each task links to the Software Dev sheet. For example, coding depends on completed design, and testing depends on code availability. According to IEC 62304, these cover the full software life cycle . | 25 | Data Collection & Curation | AI Model Development & Validati | 3 |  |  |  | 2025-10-20 | 2025-11-10 | 21 | 2025-10-20 | 2025-11-10 | 0
Electronics Hardware: Tasks include Electronic Design & Schematics (≈3 weeks) and Component Selection. Then PCB Fabrication & Assembly (≈4 weeks). After prototype arrives, Electrical Safety (IEC 60601) and EMC Pre-tests (≈2 weeks) are performed. These are in parallel with mechanical prototype builds. Deliverables (schematics, test reports) link to the Electronics sheet. The PCB build depends on completed schematics; safety testing depends on prototype readiness. | 29 | System Integration of Model | AI Model Development & Validati | 2 | Internal Verification (metrics, robustness, bias) |  |  | 2025-10-20 | 2025-11-03 | 14 | 2025-10-20 | 2025-11-03 | 0
| 32 | Schematic Design & Component Selection | Electronics Hardware Developmen | 3 |  |  |  | 2025-10-20 | 2025-11-10 | 21 | 2025-10-20 | 2025-11-10 | 0
| 40 | Conceptual Mechanical Design | Mechanical Design & Engineering | 3 |  |  |  | 2025-10-20 | 2025-11-10 | 21 | 2025-10-20 | 2025-11-10 | 0
Manufacturing Transfer: Tasks involve Develop Manufacturing Plan and DMR (≈2 weeks) once the design is frozen. Conduct Pilot Build(s) (≈4 weeks) to exercise the process. Define Process Flow/Work Instructions (≈2 weeks) and qualify any Tooling/Equipment (IQ) (≈2 weeks). Perform OQ/PQ for Production Processes (≈4 weeks total). Run Final Device Testing in Production (≈2 weeks) to verify that devices from production meet specs. Train Production Staff (≈2 weeks) and finalize the Process Validation Report (≈1 week). These tasks link to the Manufacturing sheet. Dependencies: manufacturing tasks occur after design freeze and regulatory clearance. | 45 | Design Iteration (Beta) | Mechanical Design & Engineering | 2 | Mechanical Verification (accuracy, ROM, load) |  |  | 2025-10-20 | 2025-11-03 | 14 | 2025-10-20 | 2025-11-03 | 0
| 47 | Identify Contact Materials & Classification | Biocompatibiliity & Material Sa | 1 |  |  |  | 2025-10-20 | 2025-10-27 | 7 | 2025-10-20 | 2025-10-27 | 0
| 52 | Biocompatibility Report | Biocompatibiliity & Material Sa | 2 | ISO 10993 Testing (cytotox, sensitization, irritation, etc.) |  |  | 2025-10-20 | 2025-11-03 | 14 | 2025-10-20 | 2025-11-03 | 0
| 53 | Select Sterilization Method & Worst-case Load | Sterilization Process Validatio | 1 |  |  |  | 2025-10-20 | 2025-10-27 | 7 | 2025-10-20 | 2025-10-27 | 0
| 60 | Package & Label Design | Packaging & Labeling Validation | 2 |  |  |  | 2025-10-20 | 2025-11-03 | 14 | 2025-10-20 | 2025-11-03 | 0
| 68 | Use Specification & User Profiles | Usability & Human Factors Engin | 1 |  |  |  | 2025-10-20 | 2025-10-27 | 7 | 2025-10-20 | 2025-10-27 | 0
| 75 | Threat Modeling & Cyber Risk Assessment | Cybersecurity & Data Privacy | 1 |  |  |  | 2025-10-20 | 2025-10-27 | 7 | 2025-10-20 | 2025-10-27 | 0
| 77 | Implement Security Controls | Cybersecurity & Data Privacy | 4 | Security Requirements (auth, encryption, SBOM) |  |  | 2025-10-20 | 2025-11-17 | 28 | 2025-10-20 | 2025-11-17 | 0
| 82 | Regulatory Strategy & Predicate Selection | Regulatory Submission & Documen | 1 |  |  |  | 2025-10-20 | 2025-10-27 | 7 | 2025-10-20 | 2025-10-27 | 0
| 90 | Manufacturing Plan & DMR Completion | Manufacturing Transfer & Proces | 2 |  |  |  | 2025-10-20 | 2025-11-03 | 14 | 2025-10-20 | 2025-11-03 | 0
This project management plan outlines a comprehensive pathway to develop a Class II surgical robotic navigation system from the current prototyping phase through to FDA 510(k) submission and manufacturing launch. The plan encompasses all key medical device deliverables – including design controls, risk management, biocompatibility, software and AI lifecycle, hardware development, sterilization, packaging, usability, cybersecurity, and regulatory documentation – in alignment with ISO 13485, ISO 14971, IEC 62304, IEC 60601 (and IEC 80601-2-77 for surgical robots), IEC 62366, ISO 10993, and 21 CFR Part 820 quality system requirements. Each section below describes the required tasks and deliverables, steps and tools, responsible parties, estimated timelines, and risk mitigation strategies. Tables are provided to summarize risks, potential harms, mitigations, and key deliverables for each area, and a RACI matrix assigns responsibilities across functions (Engineering, Software, QA, RA, Clinical, Manufacturing, etc.). The plan is structured to facilitate clear phase gating, traceability, and compliance with all applicable standards and regulations. | 2 | Define User Needs & System Requirements | Design Controls & Documentation | 4 | Draft Design & Development Plan | 1 |  | 2025-10-27 | 2025-11-24 | 28 | 2025-10-27 | 2025-11-24 | 0
| 11 | Preliminary Hazard Analysis (PHA) | Risk Management | 2 | Create Risk Management Plan (ISO 14971) | 10 |  | 2025-10-27 | 2025-11-10 | 14 | 2025-10-27 | 2025-11-10 | 0
| 17 | Software Requirements Specification (SRS) | Software Development | 3 | Software Development Plan (IEC 62304) | 16 |  | 2025-10-27 | 2025-11-17 | 21 | 2025-10-27 | 2025-11-17 | 0
Each task row in the Project Plan is linked (via an Excel HYPERLINK function) to the corresponding detail tab (e.g. the “Sterilization Process Validation” tasks are hyperlinked to that sheet) for full traceability. Task dependencies are indicated in the plan; for example, sterilization validation cannot occur until final device materials and packaging are defined, and usability validation follows the availability of prototypes. All durations are estimated based on typical medical device development timelines and the regulatory requirement to complete validation testing before submission . | 48 | Biocompatibility Evaluation Plan (BEP) | Biocompatibiliity & Material Sa | 2 | Identify Contact Materials & Classification | 47 |  | 2025-10-27 | 2025-11-10 | 14 | 2025-10-27 | 2025-11-10 | 0
| 54 | Sterilization Validation Plan (IQ/OQ/PQ) | Sterilization Process Validatio | 2 | Select Sterilization Method & Worst-case Load | 53 |  | 2025-10-27 | 2025-11-10 | 14 | 2025-10-27 | 2025-11-10 | 0
| 69 | Task Analysis & Critical Tasks | Usability & Human Factors Engin | 1 | Use Specification & User Profiles | 68 |  | 2025-10-27 | 2025-11-03 | 7 | 2025-10-27 | 2025-11-03 | 0
| 76 | Security Requirements (auth, encryption, SBOM) | Cybersecurity & Data Privacy | 1 | Threat Modeling & Cyber Risk Assessment | 75 |  | 2025-10-27 | 2025-11-03 | 7 | 2025-10-27 | 2025-11-03 | 0
| 83 | 510(k) Content Outline & Gap List | Regulatory Submission & Documen | 1 | Regulatory Strategy & Predicate Selection | 82 |  | 2025-10-27 | 2025-11-03 | 7 | 2025-10-27 | 2025-11-03 | 0
| 30 | Clinical/Simulated Performance Validation | AI Model Development & Validati | 4 | System Integration of Model | 29 |  | 2025-11-03 | 2025-12-01 | 28 | 2025-11-03 | 2025-12-01 | 0
| 46 | Finalize Drawings & DMR | Mechanical Design & Engineering | 2 | Design Iteration (Beta) | 45 |  | 2025-11-03 | 2025-11-17 | 14 | 2025-11-03 | 2025-11-17 | 0
| 61 | Sealing Process Development | Packaging & Labeling Validation | 1 | Package & Label Design | 60 |  | 2025-11-03 | 2025-11-10 | 7 | 2025-11-03 | 2025-11-10 | 0
| 70 | UI/Workflow Prototyping | Usability & Human Factors Engin | 2 | Task Analysis & Critical Tasks | 69 |  | 2025-11-03 | 2025-11-17 | 14 | 2025-11-03 | 2025-11-17 | 0
| 84 | Draft Device Description & Indications for Use | Regulatory Submission & Documen | 1 | 510(k) Content Outline & Gap List | 83 |  | 2025-11-03 | 2025-11-10 | 7 | 2025-11-03 | 2025-11-10 | 0
| 91 | Pilot Build (Engineering Lot) | Manufacturing Transfer & Proces | 4 | Manufacturing Plan & DMR Completion | 90 |  | 2025-11-03 | 2025-12-01 | 28 | 2025-11-03 | 2025-12-01 | 0
| 12 | System/Design FMEAs | Risk Management | 4 | Preliminary Hazard Analysis (PHA) | 11 |  | 2025-11-10 | 2025-12-08 | 28 | 2025-11-10 | 2025-12-08 | 0
| 26 | Dataset Split & Management | AI Model Development & Validati | 1 | Data Collection & Curation | 25 |  | 2025-11-10 | 2025-11-17 | 7 | 2025-11-10 | 2025-11-17 | 0
Biocompatibility: Tasks include Identify Contact Materials & Risk Assessment (≈1 week), then Select & Engage Test Lab (≈1 week). Samples are prepared from final sterilized components and sent for ISO 10993 Testing (≈8–12 weeks, including cytotoxicity, sensitization, irritation, etc.). Results are compiled into a Biocompatibility Report (≈1–2 weeks). These tasks link to the Biocompatibility sheet. Biocomp testing depends on final device/packaging (FDA expects evaluation of the finished, sterilized device ). | 33 | PCB Layout | Electronics Hardware Developmen | 3 | Schematic Design & Component Selection | 32 |  | 2025-11-10 | 2025-12-01 | 21 | 2025-11-10 | 2025-12-01 | 0
Cybersecurity & Data Privacy: Tasks begin with Threat Modeling and Cyber Risk Assessment (≈1 week) and Define Security Requirements (≈1 week). During development, Implement Security Controls (authentication, encryption, network hardening) is performed (throughout design). Later, Cybersecurity Testing (static/dynamic analysis, pen-testing, fuzz-testing) is scheduled (≈2 weeks). Remediate issues and produce an updated SBOM (≈1 week). Finally prepare Cybersecurity Documentation for submission (≈2 weeks), including a summary of controls. These link to the Cybersecurity sheet. Dependencies: security implementation follows initial design; testing follows software/hardware development. FDA now requires an SBOM of all software components , so SBOM preparation is a key deliverable. | 41 | 3D CAD Modeling & Assemblies | Mechanical Design & Engineering | 4 | Conceptual Mechanical Design | 40 |  | 2025-11-10 | 2025-12-08 | 28 | 2025-11-10 | 2025-12-08 | 0
| 49 | Select GLP Test Lab & SOW | Biocompatibiliity & Material Sa | 1 | Biocompatibility Evaluation Plan (BEP) | 48 |  | 2025-11-10 | 2025-11-17 | 7 | 2025-11-10 | 2025-11-17 | 0
| 55 | Contract Sterilizer Coordination (schedule) | Sterilization Process Validatio | 1 | Sterilization Validation Plan (IQ/OQ/PQ) | 54 |  | 2025-11-10 | 2025-11-17 | 7 | 2025-11-10 | 2025-11-17 | 0
| 62 | Packaging Validation Protocol | Packaging & Labeling Validation | 1 | Sealing Process Development | 61 |  | 2025-11-10 | 2025-11-17 | 7 | 2025-11-10 | 2025-11-17 | 0
| 85 | Compile Nonclinical Test Summaries | Regulatory Submission & Documen | 2 | Draft Device Description & Indications for Use | 84 |  | 2025-11-10 | 2025-11-24 | 14 | 2025-11-10 | 2025-11-24 | 0
| 18 | Software Architecture & Detailed Design | Software Development | 3 | Software Requirements Specification (SRS) | 17 |  | 2025-11-17 | 2025-12-08 | 21 | 2025-11-17 | 2025-12-08 | 0
AI Model Development: Tasks are Data Collection & Preparation (≈2–3 weeks), followed by Train/Develop AI Model (≈4–6 weeks) and Internal Validation (bias/robustness analysis) (≈2 weeks). After a working model, Integrate AI into System & Test (≈2 weeks) is done. If required, Clinical Performance Validation (≈4–6 weeks) follows. An Algorithm Description Document (≈1 week) is prepared for FDA submission. These tasks depend on having sufficient quality data and baseline software/hardware. The “AI Model Development” tab provides details. | 27 | Model Development & Training | AI Model Development & Validati | 5 | Dataset Split & Management | 26 |  | 2025-11-17 | 2025-12-22 | 35 | 2025-11-17 | 2025-12-22 | 0
| 50 | Prepare Final-finish Samples (incl. sterilization) | Biocompatibiliity & Material Sa | 2 | Select GLP Test Lab & SOW | 49 |  | 2025-11-17 | 2025-12-01 | 14 | 2025-11-17 | 2025-12-01 | 0
| 56 | IQ/OQ Documentation (from vendor/in-house) | Sterilization Process Validatio | 2 | Contract Sterilizer Coordination (schedule) | 55 |  | 2025-11-17 | 2025-12-01 | 14 | 2025-11-17 | 2025-12-01 | 0
| 63 | Build Test Samples (sterilized) | Packaging & Labeling Validation | 1 | Packaging Validation Protocol | 62 |  | 2025-11-17 | 2025-11-24 | 7 | 2025-11-17 | 2025-11-24 | 0
| 71 | Formative Usability Studies | Usability & Human Factors Engin | 2 | UI/Workflow Prototyping | 70 |  | 2025-11-17 | 2025-12-01 | 14 | 2025-11-17 | 2025-12-01 | 0
| 78 | Vulnerability Scans & Pen Testing | Cybersecurity & Data Privacy | 2 | Implement Security Controls | 77 |  | 2025-11-17 | 2025-12-01 | 14 | 2025-11-17 | 2025-12-01 | 0
| 3 | Develop System Architecture & Specifications | Design Controls & Documentation | 3 | Define User Needs & System Requirements | 2 |  | 2025-11-24 | 2025-12-15 | 21 | 2025-11-24 | 2025-12-15 | 0
| 64 | ISTA Transport Simulation & Integrity Tests | Packaging & Labeling Validation | 3 | Build Test Samples (sterilized) | 63 |  | 2025-11-24 | 2025-12-15 | 21 | 2025-11-24 | 2025-12-15 | 0
| 86 | Labeling/IFU Drafts | Regulatory Submission & Documen | 1 | Compile Nonclinical Test Summaries, Draft Device Description & Indications for Use | 85, 84 |  | 2025-11-24 | 2025-12-01 | 7 | 2025-11-24 | 2025-12-01 | 0
Mechanical Design: Tasks start with Conceptual Mechanical Design (≈3 weeks), then 3D CAD Modeling (≈3 weeks) of all parts. Material Selection and FEA/Tolerance Analysis* run concurrently (≈2 weeks). Next is Prototype Fabrication (Alpha build, ≈4 weeks), leading into Mechanical Verification Tests (accuracy, range, load) over ≈3 weeks. Results feed any Design Iterations (≈2 weeks). Finally Finalize Drawings & DMR (≈2 weeks). Each of these links to the Mechanical Design sheet. Dependencies: prototypes require finished CAD; testing requires built prototypes. | 31 | Algorithm Description & Documentation | AI Model Development & Validati | 1 | Clinical/Simulated Performance Validation | 30 |  | 2025-12-01 | 2025-12-08 | 7 | 2025-12-01 | 2025-12-08 | 0
| 34 | Prototype PCB Fabrication & Bring-up | Electronics Hardware Developmen | 4 | PCB Layout | 33 |  | 2025-12-01 | 2025-12-29 | 28 | 2025-12-01 | 2025-12-29 | 0
| 51 | ISO 10993 Testing (cytotox, sensitization, irritation, etc.) | Biocompatibiliity & Material Sa | 8 | Prepare Final-finish Samples (incl. sterilization) | 50 |  | 2025-12-01 | 2026-01-26 | 56 | 2025-12-01 | 2026-01-26 | 0
| 57 | PQ: Half-cycle BI Runs (3x) & Full-cycle | Sterilization Process Validatio | 4 | IQ/OQ Documentation (from vendor/in-house) | 56 |  | 2025-12-01 | 2025-12-29 | 28 | 2025-12-01 | 2025-12-29 | 0
| 72 | Summative Usability Test Protocol | Usability & Human Factors Engin | 2 | Formative Usability Studies | 71 |  | 2025-12-01 | 2025-12-15 | 14 | 2025-12-01 | 2025-12-15 | 0
| 79 | Remediation & Hardening | Cybersecurity & Data Privacy | 2 | Vulnerability Scans & Pen Testing | 78 |  | 2025-12-01 | 2025-12-15 | 14 | 2025-12-01 | 2025-12-15 | 0
| 87 | 510(k) Summary & eSTAR Assembly | Regulatory Submission & Documen | 2 | Compile Nonclinical Test Summaries, Labeling/IFU Drafts | 85, 86 |  | 2025-12-01 | 2025-12-15 | 14 | 2025-12-01 | 2025-12-15 | 0
| 92 | Define Process Flow & Work Instructions | Manufacturing Transfer & Proces | 2 | Pilot Build (Engineering Lot), Manufacturing Plan & DMR Completion | 91, 90 |  | 2025-12-01 | 2025-12-15 | 14 | 2025-12-01 | 2025-12-15 | 0
| 13 | Define Risk Controls & update requirements | Risk Management | 2 | System/Design FMEAs | 12 |  | 2025-12-08 | 2025-12-22 | 14 | 2025-12-08 | 2025-12-22 | 0
Project Plan: Tasks, Estimates, and Dependencies | 19 | Implementation (Coding) | Software Development | 8 | Software Architecture & Detailed Design | 18 |  | 2025-12-08 | 2026-02-02 | 56 | 2025-12-08 | 2026-02-02 | 0
| 42 | FEA & Tolerance Analysis | Mechanical Design & Engineering | 2 | 3D CAD Modeling & Assemblies | 41 |  | 2025-12-08 | 2025-12-22 | 14 | 2025-12-08 | 2025-12-22 | 0
| 4 | Plan Design Reviews (PDR, CDR) | Design Controls & Documentation | 1 | Develop System Architecture & Specifications | 3 |  | 2025-12-15 | 2025-12-22 | 7 | 2025-12-15 | 2025-12-22 | 0
| 5 | Create Verification Plan (trace to requirements) | Design Controls & Documentation | 2 | Plan Design Reviews (PDR, CDR), Develop System Architecture & Specifications | 3 |  | 2025-12-15 | 2025-12-29 | 14 | 2025-12-15 | 2025-12-29 | 0
| 65 | Accelerated Aging (if shelf-life claimed) | Packaging & Labeling Validation | 12 | ISTA Transport Simulation & Integrity Tests, Build Test Samples (sterilized) | 64, 63 |  | 2025-12-15 | 2026-03-09 | 84 | 2025-12-15 | 2026-03-09 | 0
| 73 | Summative Usability Test Execution | Usability & Human Factors Engin | 2 | Summative Usability Test Protocol | 72 |  | 2025-12-15 | 2025-12-29 | 14 | 2025-12-15 | 2025-12-29 | 0
| 80 | SBOM Creation | Cybersecurity & Data Privacy | 1 | Remediation & Hardening | 79 |  | 2025-12-15 | 2025-12-22 | 7 | 2025-12-15 | 2025-12-22 | 0
| 88 | Internal Submission Review & Approval | Regulatory Submission & Documen | 1 | 510(k) Summary & eSTAR Assembly | 87 |  | 2025-12-15 | 2025-12-22 | 7 | 2025-12-15 | 2025-12-22 | 0
| 93 | Tooling/Equipment IQ | Manufacturing Transfer & Proces | 2 | Define Process Flow & Work Instructions | 92 |  | 2025-12-15 | 2025-12-29 | 14 | 2025-12-15 | 2025-12-29 | 0
Traceability of Deliverables Spreadsheet Analysis | 14 | Verify Risk Controls Implemented (tests) | Risk Management | 2 | Define Risk Controls & update requirements | 13 |  | 2025-12-22 | 2026-01-05 | 14 | 2025-12-22 | 2026-01-05 | 0
| 28 | Internal Verification (metrics, robustness, bias) | AI Model Development & Validati | 2 | Model Development & Training | 27 |  | 2025-12-22 | 2026-01-05 | 14 | 2025-12-22 | 2026-01-05 | 0
Regulatory Submission: Tasks include Compile Technical Documentation (ongoing) and write the Device Description, Indications for Use, and Predicate Comparison (≈1 week). As tests complete, Summarize Nonclinical Testing Results by category (mechanical, software, safety, etc.) (≈2 weeks). Develop Labeling and IFU drafts (≈1 week). Write the 510(k) Summary and Assemble Submission (≈3–4 weeks in months 8–9) . Conduct internal Submission Review/Approval (≈1 week), then Submit via FDA’s eSTAR portal (≈1 day). These link to the Regulatory sheet. All tasks depend on completed test reports and design outputs. | 43 | Prototype Fabrication (Alpha) | Mechanical Design & Engineering | 4 | FEA & Tolerance Analysis, 3D CAD Modeling & Assemblies | 42, 41 |  | 2025-12-22 | 2026-01-19 | 28 | 2025-12-22 | 2026-01-19 | 0
| 81 | Cybersecurity Documentation (FDA) | Cybersecurity & Data Privacy | 2 | SBOM Creation | 80 |  | 2025-12-22 | 2026-01-05 | 14 | 2025-12-22 | 2026-01-05 | 0
| 89 | Submit 510(k) | Regulatory Submission & Documen | 0.2 | Internal Submission Review & Approval | 88 |  | 2025-12-22 | 2025-12-24 | 2 | 2025-12-22 | 2025-12-24 | 0
| 7 | Execute Design Verification (bench/system) | Design Controls & Documentation | 6 | Generate Design Outputs (CAD, Schematics, Code baselines), Create Verification Plan (trace to requirements) | 5 |  | 2025-12-29 | 2026-02-09 | 42 | 2025-12-29 | 2026-02-09 | 0
Sterilization Validation: Tasks begin with Determine Sterilization Method & Needs (≈1 week), followed by Develop Sterilization Validation Plan (≈2 weeks). Engage a sterilization provider and perform Equipment IQ/OQ (≈2 weeks) and then Performance Qualification (PQ) (≈4 weeks, including half-cycle BI tests and full-cycle validations). Finally, compile Sterilization Validation Report (≈1 week). Each task links to the Sterilization sheet. These depend on finalized packaging and prototypes. As ISO 11135 emphasizes, sterilization validation comprises IQ, OQ, and PQ stages . | 35 | Bench/Engineering Tests | Electronics Hardware Developmen | 2 | Prototype PCB Fabrication & Bring-up | 34 |  | 2025-12-29 | 2026-01-12 | 14 | 2025-12-29 | 2026-01-12 | 0
| 58 | EO Residuals / Physical PQ & Functional Checks | Sterilization Process Validatio | 2 | PQ: Half-cycle BI Runs (3x) & Full-cycle | 57 |  | 2025-12-29 | 2026-01-12 | 14 | 2025-12-29 | 2026-01-12 | 0
| 74 | Usability Validation Report | Usability & Human Factors Engin | 1 | Summative Usability Test Execution | 73 |  | 2025-12-29 | 2026-01-05 | 7 | 2025-12-29 | 2026-01-05 | 0
| 94 | Process OQ | Manufacturing Transfer & Proces | 2 | Tooling/Equipment IQ | 93 |  | 2025-12-29 | 2026-01-12 | 14 | 2025-12-29 | 2026-01-12 | 0
| 15 | Residual Risk Evaluation & Risk Report | Risk Management | 1 | Verify Risk Controls Implemented (tests) | 14 |  | 2026-01-05 | 2026-01-12 | 7 | 2026-01-05 | 2026-01-12 | 0
| 36 | EMC Pre-compliance Testing | Electronics Hardware Developmen | 1 | Bench/Engineering Tests | 35 |  | 2026-01-12 | 2026-01-19 | 7 | 2026-01-12 | 2026-01-19 | 0
| 59 | Sterilization Validation Report | Sterilization Process Validatio | 1 | EO Residuals / Physical PQ & Functional Checks | 58 |  | 2026-01-12 | 2026-01-19 | 7 | 2026-01-12 | 2026-01-19 | 0
| 95 | Process PQ (3 lots) | Manufacturing Transfer & Proces | 4 | Process OQ | 94 |  | 2026-01-12 | 2026-02-09 | 28 | 2026-01-12 | 2026-02-09 | 0
Packaging & Labeling: Tasks include Package and Label Design (≈2 weeks) and drafting Packaging Validation Protocol (≈1 week). Then Prepare Test Samples (sterilize and assemble, ≈1 week). Send for Transport Simulation Testing (ISTA) and Seal Integrity Tests (≈3 weeks total). If shelf-life is claimed, Accelerated Aging is run (≈12 weeks). Label Adhesion/Legibility Testing occurs under environmental stress (≈4 weeks). Finally, write the Packaging Validation Report (≈1 week). These link to the Packaging sheet. Dependencies: packaging design requires final device dimensions; tests depend on assembled sterile units. FDA QSR requires rigorous labeling control (legibility and attachment) , so label-related tasks reflect those regulations. | 37 | 60601 Safety Design Review | Electronics Hardware Developmen | 1 | EMC Pre-compliance Testing, Bench/Engineering Tests | 36, 35 |  | 2026-01-19 | 2026-01-26 | 7 | 2026-01-19 | 2026-01-26 | 0
| 44 | Mechanical Verification (accuracy, ROM, load) | Mechanical Design & Engineering | 3 | Prototype Fabrication (Alpha) | 43 |  | 2026-01-19 | 2026-02-09 | 21 | 2026-01-19 | 2026-02-09 | 0
| 38 | Design Freeze (Electronics) | Electronics Hardware Developmen | 1 | EMC Pre-compliance Testing, 60601 Safety Design Review | 36, 37 |  | 2026-01-26 | 2026-02-02 | 7 | 2026-01-26 | 2026-02-02 | 0
| 20 | Unit Testing | Software Development | 2 | Implementation (Coding) | 19 |  | 2026-02-02 | 2026-02-16 | 14 | 2026-02-02 | 2026-02-16 | 0
Usability & Human Factors: Tasks include Define User Profiles/Use Environment (≈1 week) and Task Analysis (≈1 week). Based on risk analysis, design a User Interface Prototype (≈2 weeks) and conduct Formative Usability Tests (≈2 weeks) to identify use errors. Then plan Summative Usability (Validation) Test (≈2 weeks) and execute it (≈2 weeks), analyzing results in a Usability Report (≈1 week). These tasks link to the Usability sheet. Dependencies: HF tasks depend on the availability of functional prototypes (hardware+software). Insights from HF testing feed design updates. FDA guidance stresses including usability data as part of validation . | 39 | Formal 60601/EMC/80601 Testing (3rd party) | Electronics Hardware Developmen | 3 | Design Freeze (Electronics) | 38 |  | 2026-02-02 | 2026-02-23 | 21 | 2026-02-02 | 2026-02-23 | 0
| 8 | Design Validation (simulated use) | Design Controls & Documentation | 4 | Execute Design Verification (bench/system) | 7 |  | 2026-02-09 | 2026-03-09 | 28 | 2026-02-09 | 2026-03-09 | 0
| 96 | Train Production Personnel | Manufacturing Transfer & Proces | 2 | Process PQ (3 lots), Define Process Flow & Work Instructions | 95, 92 |  | 2026-02-09 | 2026-02-23 | 14 | 2026-02-09 | 2026-02-23 | 0
Design Controls & Documentation: Key tasks include Draft Design & Development Plan (≈2–3 weeks) once project scope is set; Gather User Needs and Define Requirements (≈3–4 weeks, depends on initial concept documents); and Develop System Architecture & Specifications (≈2 weeks). Design Reviews and Verification planning occur at milestones. A Validation (simulated use) study (≈3–4 weeks) is planned near design freeze. Each task in the plan is hyperlinked to the “Design Controls & Documentation” tab for details. These depend on having a finalized concept and requirements baseline . | 21 | Integration Testing (HW/SW-in-the-loop) | Software Development | 3 | Unit Testing | 20 |  | 2026-02-16 | 2026-03-09 | 21 | 2026-02-16 | 2026-03-09 | 0
| 97 | Process Validation Report | Manufacturing Transfer & Proces | 1 | Train Production Personnel, Process PQ (3 lots) | 96, 95 |  | 2026-02-23 | 2026-03-02 | 7 | 2026-02-23 | 2026-03-02 | 0
Description & Standards: Establish robust design control processes to ensure the device meets user needs, is safe/effective, and complies with regulatory requirements . Under ISO 13485 and FDA QSR, we must maintain a Design and Development Plan defining stages, reviews, verification/validation, responsibilities, and traceability . Key design control elements include documenting user needs, design inputs (requirements), design outputs (specifications, drawings, code, etc.), design reviews, design verification (does output meet inputs?), design validation (does device meet user needs in intended use), and design transfer to manufacturing . All activities and decisions are recorded in the Design History File (DHF) . | 9 | Design Transfer & DMR Compilation | Design Controls & Documentation | 4 | Design Validation (simulated use) | 8 |  | 2026-03-09 | 2026-04-06 | 28 | 2026-03-09 | 2026-04-06 | 0
| 22 | Software Verification (System Tests) | Software Development | 4 | Integration Testing (HW/SW-in-the-loop) | 21 |  | 2026-03-09 | 2026-04-06 | 28 | 2026-03-09 | 2026-04-06 | 0
| 66 | Label Legibility/Adhesion Testing | Packaging & Labeling Validation | 2 | Accelerated Aging (if shelf-life claimed), Package & Label Design | 65, 60 |  | 2026-03-09 | 2026-03-23 | 14 | 2026-03-09 | 2026-03-23 | 0
| 67 | Packaging Validation Report | Packaging & Labeling Validation | 1 | Label Legibility/Adhesion Testing, ISTA Transport Simulation & Integrity Tests | 66, 64 |  | 2026-03-23 | 2026-03-30 | 7 | 2026-03-23 | 2026-03-30 | 0
Risk Management: Project Plan lists Perform Preliminary Hazard Analysis (≈1 week) early on, and iterative Risk Analysis Updates (≈2 weeks each round) after design iterations. These tasks link to the Risk Management sheet. They depend on draft designs (hardware, software, use scenarios). The risk file is continually updated and then summarized in the submission. | 23 | Software Validation (simulated use) | Software Development | 3 | Software Verification (System Tests) | 22 |  | 2026-04-06 | 2026-04-27 | 21 | 2026-04-06 | 2026-04-27 | 0
| 24 | Release & Configuration Audit | Software Development | 1 | Software Validation (simulated use) | 23 |  | 2026-04-27 | 2026-05-04 | 7 | 2026-04-27 | 2026-05-04 | 0

## Gantt Data
ID | Task Name | Area (Tab) | Responsible Party | Start Date | Finish Date | Duration (days) | Depends On (IDs)
1 | Draft Design & Development Plan | Design Controls & Documentation |  | 2025-10-20 | 2025-10-27 | 7
2 | Define User Needs & System Requirements | Design Controls & Documentation |  | 2025-10-27 | 2025-11-24 | 28 | 1
3 | Develop System Architecture & Specifications | Design Controls & Documentation |  | 2025-11-24 | 2025-12-15 | 21 | 2
4 | Plan Design Reviews (PDR, CDR) | Design Controls & Documentation |  | 2025-12-15 | 2025-12-22 | 7 | 3
5 | Create Verification Plan (trace to requirements) | Design Controls & Documentation |  | 2025-12-15 | 2025-12-29 | 14 | 3
6 | Generate Design Outputs (CAD, Schematics, Code baselines) | Design Controls & Documentation |  | 2025-10-20 | 2025-12-15 | 56 | 2025-05-03
7 | Execute Design Verification (bench/system) | Design Controls & Documentation |  | 2025-12-29 | 2026-02-09 | 42 | 5
8 | Design Validation (simulated use) | Design Controls & Documentation |  | 2026-02-09 | 2026-03-09 | 28 | 7
9 | Design Transfer & DMR Compilation | Design Controls & Documentation |  | 2026-03-09 | 2026-04-06 | 28 | 8
10 | Create Risk Management Plan (ISO 14971) | Risk Management |  | 2025-10-20 | 2025-10-27 | 7
11 | Preliminary Hazard Analysis (PHA) | Risk Management |  | 2025-10-27 | 2025-11-10 | 14 | 10
12 | System/Design FMEAs | Risk Management |  | 2025-11-10 | 2025-12-08 | 28 | 11
13 | Define Risk Controls & update requirements | Risk Management |  | 2025-12-08 | 2025-12-22 | 14 | 12
14 | Verify Risk Controls Implemented (tests) | Risk Management |  | 2025-12-22 | 2026-01-05 | 14 | 13
15 | Residual Risk Evaluation & Risk Report | Risk Management |  | 2026-01-05 | 2026-01-12 | 7 | 14
16 | Software Development Plan (IEC 62304) | Software Development |  | 2025-10-20 | 2025-10-27 | 7
17 | Software Requirements Specification (SRS) | Software Development |  | 2025-10-27 | 2025-11-17 | 21 | 16
18 | Software Architecture & Detailed Design | Software Development |  | 2025-11-17 | 2025-12-08 | 21 | 17
19 | Implementation (Coding) | Software Development |  | 2025-12-08 | 2026-02-02 | 56 | 18
20 | Unit Testing | Software Development |  | 2026-02-02 | 2026-02-16 | 14 | 19
21 | Integration Testing (HW/SW-in-the-loop) | Software Development |  | 2026-02-16 | 2026-03-09 | 21 | 20
22 | Software Verification (System Tests) | Software Development |  | 2026-03-09 | 2026-04-06 | 28 | 21
23 | Software Validation (simulated use) | Software Development |  | 2026-04-06 | 2026-04-27 | 21 | 22
24 | Release & Configuration Audit | Software Development |  | 2026-04-27 | 2026-05-04 | 7 | 23
25 | Data Collection & Curation | AI Model Development & Validati |  | 2025-10-20 | 2025-11-10 | 21
26 | Dataset Split & Management | AI Model Development & Validati |  | 2025-11-10 | 2025-11-17 | 7 | 25
27 | Model Development & Training | AI Model Development & Validati |  | 2025-11-17 | 2025-12-22 | 35 | 26
28 | Internal Verification (metrics, robustness, bias) | AI Model Development & Validati |  | 2025-12-22 | 2026-01-05 | 14 | 27
29 | System Integration of Model | AI Model Development & Validati |  | 2025-10-20 | 2025-11-03 | 14
30 | Clinical/Simulated Performance Validation | AI Model Development & Validati |  | 2025-11-03 | 2025-12-01 | 28 | 29
31 | Algorithm Description & Documentation | AI Model Development & Validati |  | 2025-12-01 | 2025-12-08 | 7 | 30
32 | Schematic Design & Component Selection | Electronics Hardware Developmen |  | 2025-10-20 | 2025-11-10 | 21
33 | PCB Layout | Electronics Hardware Developmen |  | 2025-11-10 | 2025-12-01 | 21 | 32
34 | Prototype PCB Fabrication & Bring-up | Electronics Hardware Developmen |  | 2025-12-01 | 2025-12-29 | 28 | 33
35 | Bench/Engineering Tests | Electronics Hardware Developmen |  | 2025-12-29 | 2026-01-12 | 14 | 34
36 | EMC Pre-compliance Testing | Electronics Hardware Developmen |  | 2026-01-12 | 2026-01-19 | 7 | 35
37 | 60601 Safety Design Review | Electronics Hardware Developmen |  | 2026-01-19 | 2026-01-26 | 7 | 36, 35
38 | Design Freeze (Electronics) | Electronics Hardware Developmen |  | 2026-01-26 | 2026-02-02 | 7 | 36, 37
39 | Formal 60601/EMC/80601 Testing (3rd party) | Electronics Hardware Developmen |  | 2026-02-02 | 2026-02-23 | 21 | 38
40 | Conceptual Mechanical Design | Mechanical Design & Engineering |  | 2025-10-20 | 2025-11-10 | 21
41 | 3D CAD Modeling & Assemblies | Mechanical Design & Engineering |  | 2025-11-10 | 2025-12-08 | 28 | 40
42 | FEA & Tolerance Analysis | Mechanical Design & Engineering |  | 2025-12-08 | 2025-12-22 | 14 | 41
43 | Prototype Fabrication (Alpha) | Mechanical Design & Engineering |  | 2025-12-22 | 2026-01-19 | 28 | 42, 41
44 | Mechanical Verification (accuracy, ROM, load) | Mechanical Design & Engineering |  | 2026-01-19 | 2026-02-09 | 21 | 43
45 | Design Iteration (Beta) | Mechanical Design & Engineering |  | 2025-10-20 | 2025-11-03 | 14
46 | Finalize Drawings & DMR | Mechanical Design & Engineering |  | 2025-11-03 | 2025-11-17 | 14 | 45
47 | Identify Contact Materials & Classification | Biocompatibiliity & Material Sa |  | 2025-10-20 | 2025-10-27 | 7
48 | Biocompatibility Evaluation Plan (BEP) | Biocompatibiliity & Material Sa |  | 2025-10-27 | 2025-11-10 | 14 | 47
49 | Select GLP Test Lab & SOW | Biocompatibiliity & Material Sa |  | 2025-11-10 | 2025-11-17 | 7 | 48
50 | Prepare Final-finish Samples (incl. sterilization) | Biocompatibiliity & Material Sa |  | 2025-11-17 | 2025-12-01 | 14 | 49
51 | ISO 10993 Testing (cytotox, sensitization, irritation, etc.) | Biocompatibiliity & Material Sa |  | 2025-12-01 | 2026-01-26 | 56 | 50
52 | Biocompatibility Report | Biocompatibiliity & Material Sa |  | 2025-10-20 | 2025-11-03 | 14
53 | Select Sterilization Method & Worst-case Load | Sterilization Process Validatio |  | 2025-10-20 | 2025-10-27 | 7
54 | Sterilization Validation Plan (IQ/OQ/PQ) | Sterilization Process Validatio |  | 2025-10-27 | 2025-11-10 | 14 | 53
55 | Contract Sterilizer Coordination (schedule) | Sterilization Process Validatio |  | 2025-11-10 | 2025-11-17 | 7 | 54
56 | IQ/OQ Documentation (from vendor/in-house) | Sterilization Process Validatio |  | 2025-11-17 | 2025-12-01 | 14 | 55
57 | PQ: Half-cycle BI Runs (3x) & Full-cycle | Sterilization Process Validatio |  | 2025-12-01 | 2025-12-29 | 28 | 56
58 | EO Residuals / Physical PQ & Functional Checks | Sterilization Process Validatio |  | 2025-12-29 | 2026-01-12 | 14 | 57
59 | Sterilization Validation Report | Sterilization Process Validatio |  | 2026-01-12 | 2026-01-19 | 7 | 58
60 | Package & Label Design | Packaging & Labeling Validation |  | 2025-10-20 | 2025-11-03 | 14
61 | Sealing Process Development | Packaging & Labeling Validation |  | 2025-11-03 | 2025-11-10 | 7 | 60
62 | Packaging Validation Protocol | Packaging & Labeling Validation |  | 2025-11-10 | 2025-11-17 | 7 | 61
63 | Build Test Samples (sterilized) | Packaging & Labeling Validation |  | 2025-11-17 | 2025-11-24 | 7 | 62
64 | ISTA Transport Simulation & Integrity Tests | Packaging & Labeling Validation |  | 2025-11-24 | 2025-12-15 | 21 | 63
65 | Accelerated Aging (if shelf-life claimed) | Packaging & Labeling Validation |  | 2025-12-15 | 2026-03-09 | 84 | 64, 63
66 | Label Legibility/Adhesion Testing | Packaging & Labeling Validation |  | 2026-03-09 | 2026-03-23 | 14 | 65, 60
67 | Packaging Validation Report | Packaging & Labeling Validation |  | 2026-03-23 | 2026-03-30 | 7 | 66, 64
68 | Use Specification & User Profiles | Usability & Human Factors Engin |  | 2025-10-20 | 2025-10-27 | 7
69 | Task Analysis & Critical Tasks | Usability & Human Factors Engin |  | 2025-10-27 | 2025-11-03 | 7 | 68
70 | UI/Workflow Prototyping | Usability & Human Factors Engin |  | 2025-11-03 | 2025-11-17 | 14 | 69
71 | Formative Usability Studies | Usability & Human Factors Engin |  | 2025-11-17 | 2025-12-01 | 14 | 70
72 | Summative Usability Test Protocol | Usability & Human Factors Engin |  | 2025-12-01 | 2025-12-15 | 14 | 71
73 | Summative Usability Test Execution | Usability & Human Factors Engin |  | 2025-12-15 | 2025-12-29 | 14 | 72
74 | Usability Validation Report | Usability & Human Factors Engin |  | 2025-12-29 | 2026-01-05 | 7 | 73
75 | Threat Modeling & Cyber Risk Assessment | Cybersecurity & Data Privacy |  | 2025-10-20 | 2025-10-27 | 7
76 | Security Requirements (auth, encryption, SBOM) | Cybersecurity & Data Privacy |  | 2025-10-27 | 2025-11-03 | 7 | 75
77 | Implement Security Controls | Cybersecurity & Data Privacy |  | 2025-10-20 | 2025-11-17 | 28
78 | Vulnerability Scans & Pen Testing | Cybersecurity & Data Privacy |  | 2025-11-17 | 2025-12-01 | 14 | 77
79 | Remediation & Hardening | Cybersecurity & Data Privacy |  | 2025-12-01 | 2025-12-15 | 14 | 78
80 | SBOM Creation | Cybersecurity & Data Privacy |  | 2025-12-15 | 2025-12-22 | 7 | 79
81 | Cybersecurity Documentation (FDA) | Cybersecurity & Data Privacy |  | 2025-12-22 | 2026-01-05 | 14 | 80
82 | Regulatory Strategy & Predicate Selection | Regulatory Submission & Documen |  | 2025-10-20 | 2025-10-27 | 7
83 | 510(k) Content Outline & Gap List | Regulatory Submission & Documen |  | 2025-10-27 | 2025-11-03 | 7 | 82
84 | Draft Device Description & Indications for Use | Regulatory Submission & Documen |  | 2025-11-03 | 2025-11-10 | 7 | 83
85 | Compile Nonclinical Test Summaries | Regulatory Submission & Documen |  | 2025-11-10 | 2025-11-24 | 14 | 84
86 | Labeling/IFU Drafts | Regulatory Submission & Documen |  | 2025-11-24 | 2025-12-01 | 7 | 85, 84
87 | 510(k) Summary & eSTAR Assembly | Regulatory Submission & Documen |  | 2025-12-01 | 2025-12-15 | 14 | 85, 86
88 | Internal Submission Review & Approval | Regulatory Submission & Documen |  | 2025-12-15 | 2025-12-22 | 7 | 87
89 | Submit 510(k) | Regulatory Submission & Documen |  | 2025-12-22 | 2025-12-24 | 2 | 88
90 | Manufacturing Plan & DMR Completion | Manufacturing Transfer & Proces |  | 2025-10-20 | 2025-11-03 | 14
91 | Pilot Build (Engineering Lot) | Manufacturing Transfer & Proces |  | 2025-11-03 | 2025-12-01 | 28 | 90
92 | Define Process Flow & Work Instructions | Manufacturing Transfer & Proces |  | 2025-12-01 | 2025-12-15 | 14 | 91, 90
93 | Tooling/Equipment IQ | Manufacturing Transfer & Proces |  | 2025-12-15 | 2025-12-29 | 14 | 92
94 | Process OQ | Manufacturing Transfer & Proces |  | 2025-12-29 | 2026-01-12 | 14 | 93
95 | Process PQ (3 lots) | Manufacturing Transfer & Proces |  | 2026-01-12 | 2026-02-09 | 28 | 94
96 | Train Production Personnel | Manufacturing Transfer & Proces |  | 2026-02-09 | 2026-02-23 | 14 | 95, 92
97 | Process Validation Report | Manufacturing Transfer & Proces |  | 2026-02-23 | 2026-03-02 | 7 | 96, 95

## Gantt Data (Leveled)
ID | Task Name | Area (Tab) | Responsible Party | Start Date | Finish Date | Duration (days) | Depends On (IDs)
1 | Draft Design & Development Plan | Design Controls & Documentation |  | 2025-10-20 | 2025-10-27 | 7
2 | Define User Needs & System Requirements | Design Controls & Documentation |  | 2025-10-27 | 2025-11-24 | 28 | 1
3 | Develop System Architecture & Specifications | Design Controls & Documentation |  | 2025-11-24 | 2025-12-15 | 21 | 2
4 | Plan Design Reviews (PDR, CDR) | Design Controls & Documentation |  | 2025-12-15 | 2025-12-22 | 7 | 3
5 | Create Verification Plan (trace to requirements) | Design Controls & Documentation |  | 2025-12-15 | 2025-12-29 | 14 | 3
6 | Generate Design Outputs (CAD, Schematics, Code baselines) | Design Controls & Documentation |  | 2025-10-20 | 2025-12-15 | 56 | 2025-05-03
7 | Execute Design Verification (bench/system) | Design Controls & Documentation |  | 2025-12-29 | 2026-02-09 | 42 | 5
8 | Design Validation (simulated use) | Design Controls & Documentation |  | 2026-02-09 | 2026-03-09 | 28 | 7
9 | Design Transfer & DMR Compilation | Design Controls & Documentation |  | 2026-03-09 | 2026-04-06 | 28 | 8
10 | Create Risk Management Plan (ISO 14971) | Risk Management |  | 2025-10-20 | 2025-10-27 | 7
11 | Preliminary Hazard Analysis (PHA) | Risk Management |  | 2025-10-27 | 2025-11-10 | 14 | 10
12 | System/Design FMEAs | Risk Management |  | 2025-11-10 | 2025-12-08 | 28 | 11
13 | Define Risk Controls & update requirements | Risk Management |  | 2025-12-08 | 2025-12-22 | 14 | 12
14 | Verify Risk Controls Implemented (tests) | Risk Management |  | 2025-12-22 | 2026-01-05 | 14 | 13
15 | Residual Risk Evaluation & Risk Report | Risk Management |  | 2026-01-05 | 2026-01-12 | 7 | 14
16 | Software Development Plan (IEC 62304) | Software Development |  | 2025-10-20 | 2025-10-27 | 7
17 | Software Requirements Specification (SRS) | Software Development |  | 2025-10-27 | 2025-11-17 | 21 | 16
18 | Software Architecture & Detailed Design | Software Development |  | 2025-11-17 | 2025-12-08 | 21 | 17
19 | Implementation (Coding) | Software Development |  | 2025-12-08 | 2026-02-02 | 56 | 18
20 | Unit Testing | Software Development |  | 2026-02-02 | 2026-02-16 | 14 | 19
21 | Integration Testing (HW/SW-in-the-loop) | Software Development |  | 2026-02-16 | 2026-03-09 | 21 | 20
22 | Software Verification (System Tests) | Software Development |  | 2026-03-09 | 2026-04-06 | 28 | 21
23 | Software Validation (simulated use) | Software Development |  | 2026-04-06 | 2026-04-27 | 21 | 22
24 | Release & Configuration Audit | Software Development |  | 2026-04-27 | 2026-05-04 | 7 | 23
25 | Data Collection & Curation | AI Model Development & Validati |  | 2025-10-20 | 2025-11-10 | 21
26 | Dataset Split & Management | AI Model Development & Validati |  | 2025-11-10 | 2025-11-17 | 7 | 25
27 | Model Development & Training | AI Model Development & Validati |  | 2025-11-17 | 2025-12-22 | 35 | 26
28 | Internal Verification (metrics, robustness, bias) | AI Model Development & Validati |  | 2025-12-22 | 2026-01-05 | 14 | 27
29 | System Integration of Model | AI Model Development & Validati |  | 2025-10-20 | 2025-11-03 | 14
30 | Clinical/Simulated Performance Validation | AI Model Development & Validati |  | 2025-11-03 | 2025-12-01 | 28 | 29
31 | Algorithm Description & Documentation | AI Model Development & Validati |  | 2025-12-01 | 2025-12-08 | 7 | 30
32 | Schematic Design & Component Selection | Electronics Hardware Developmen |  | 2025-10-20 | 2025-11-10 | 21
33 | PCB Layout | Electronics Hardware Developmen |  | 2025-11-10 | 2025-12-01 | 21 | 32
34 | Prototype PCB Fabrication & Bring-up | Electronics Hardware Developmen |  | 2025-12-01 | 2025-12-29 | 28 | 33
35 | Bench/Engineering Tests | Electronics Hardware Developmen |  | 2025-12-29 | 2026-01-12 | 14 | 34
36 | EMC Pre-compliance Testing | Electronics Hardware Developmen |  | 2026-01-12 | 2026-01-19 | 7 | 35
37 | 60601 Safety Design Review | Electronics Hardware Developmen |  | 2026-01-19 | 2026-01-26 | 7 | 36, 35
38 | Design Freeze (Electronics) | Electronics Hardware Developmen |  | 2026-01-26 | 2026-02-02 | 7 | 36, 37
39 | Formal 60601/EMC/80601 Testing (3rd party) | Electronics Hardware Developmen |  | 2026-02-02 | 2026-02-23 | 21 | 38
40 | Conceptual Mechanical Design | Mechanical Design & Engineering |  | 2025-10-20 | 2025-11-10 | 21
41 | 3D CAD Modeling & Assemblies | Mechanical Design & Engineering |  | 2025-11-10 | 2025-12-08 | 28 | 40
42 | FEA & Tolerance Analysis | Mechanical Design & Engineering |  | 2025-12-08 | 2025-12-22 | 14 | 41
43 | Prototype Fabrication (Alpha) | Mechanical Design & Engineering |  | 2025-12-22 | 2026-01-19 | 28 | 42, 41
44 | Mechanical Verification (accuracy, ROM, load) | Mechanical Design & Engineering |  | 2026-01-19 | 2026-02-09 | 21 | 43
45 | Design Iteration (Beta) | Mechanical Design & Engineering |  | 2025-10-20 | 2025-11-03 | 14
46 | Finalize Drawings & DMR | Mechanical Design & Engineering |  | 2025-11-03 | 2025-11-17 | 14 | 45
47 | Identify Contact Materials & Classification | Biocompatibiliity & Material Sa |  | 2025-10-20 | 2025-10-27 | 7
48 | Biocompatibility Evaluation Plan (BEP) | Biocompatibiliity & Material Sa |  | 2025-10-27 | 2025-11-10 | 14 | 47
49 | Select GLP Test Lab & SOW | Biocompatibiliity & Material Sa |  | 2025-11-10 | 2025-11-17 | 7 | 48
50 | Prepare Final-finish Samples (incl. sterilization) | Biocompatibiliity & Material Sa |  | 2025-11-17 | 2025-12-01 | 14 | 49
51 | ISO 10993 Testing (cytotox, sensitization, irritation, etc.) | Biocompatibiliity & Material Sa |  | 2025-12-01 | 2026-01-26 | 56 | 50
52 | Biocompatibility Report | Biocompatibiliity & Material Sa |  | 2025-10-20 | 2025-11-03 | 14
53 | Select Sterilization Method & Worst-case Load | Sterilization Process Validatio |  | 2025-10-20 | 2025-10-27 | 7
54 | Sterilization Validation Plan (IQ/OQ/PQ) | Sterilization Process Validatio |  | 2025-10-27 | 2025-11-10 | 14 | 53
55 | Contract Sterilizer Coordination (schedule) | Sterilization Process Validatio |  | 2025-11-10 | 2025-11-17 | 7 | 54
56 | IQ/OQ Documentation (from vendor/in-house) | Sterilization Process Validatio |  | 2025-11-17 | 2025-12-01 | 14 | 55
57 | PQ: Half-cycle BI Runs (3x) & Full-cycle | Sterilization Process Validatio |  | 2025-12-01 | 2025-12-29 | 28 | 56
58 | EO Residuals / Physical PQ & Functional Checks | Sterilization Process Validatio |  | 2025-12-29 | 2026-01-12 | 14 | 57
59 | Sterilization Validation Report | Sterilization Process Validatio |  | 2026-01-12 | 2026-01-19 | 7 | 58
60 | Package & Label Design | Packaging & Labeling Validation |  | 2025-10-20 | 2025-11-03 | 14
61 | Sealing Process Development | Packaging & Labeling Validation |  | 2025-11-03 | 2025-11-10 | 7 | 60
62 | Packaging Validation Protocol | Packaging & Labeling Validation |  | 2025-11-10 | 2025-11-17 | 7 | 61
63 | Build Test Samples (sterilized) | Packaging & Labeling Validation |  | 2025-11-17 | 2025-11-24 | 7 | 62
64 | ISTA Transport Simulation & Integrity Tests | Packaging & Labeling Validation |  | 2025-11-24 | 2025-12-15 | 21 | 63
65 | Accelerated Aging (if shelf-life claimed) | Packaging & Labeling Validation |  | 2025-12-15 | 2026-03-09 | 84 | 64, 63
66 | Label Legibility/Adhesion Testing | Packaging & Labeling Validation |  | 2026-03-09 | 2026-03-23 | 14 | 65, 60
67 | Packaging Validation Report | Packaging & Labeling Validation |  | 2026-03-23 | 2026-03-30 | 7 | 66, 64
68 | Use Specification & User Profiles | Usability & Human Factors Engin |  | 2025-10-20 | 2025-10-27 | 7
69 | Task Analysis & Critical Tasks | Usability & Human Factors Engin |  | 2025-10-27 | 2025-11-03 | 7 | 68
70 | UI/Workflow Prototyping | Usability & Human Factors Engin |  | 2025-11-03 | 2025-11-17 | 14 | 69
71 | Formative Usability Studies | Usability & Human Factors Engin |  | 2025-11-17 | 2025-12-01 | 14 | 70
72 | Summative Usability Test Protocol | Usability & Human Factors Engin |  | 2025-12-01 | 2025-12-15 | 14 | 71
73 | Summative Usability Test Execution | Usability & Human Factors Engin |  | 2025-12-15 | 2025-12-29 | 14 | 72
74 | Usability Validation Report | Usability & Human Factors Engin |  | 2025-12-29 | 2026-01-05 | 7 | 73
75 | Threat Modeling & Cyber Risk Assessment | Cybersecurity & Data Privacy |  | 2025-10-20 | 2025-10-27 | 7
76 | Security Requirements (auth, encryption, SBOM) | Cybersecurity & Data Privacy |  | 2025-10-27 | 2025-11-03 | 7 | 75
77 | Implement Security Controls | Cybersecurity & Data Privacy |  | 2025-10-20 | 2025-11-17 | 28
78 | Vulnerability Scans & Pen Testing | Cybersecurity & Data Privacy |  | 2025-11-17 | 2025-12-01 | 14 | 77
79 | Remediation & Hardening | Cybersecurity & Data Privacy |  | 2025-12-01 | 2025-12-15 | 14 | 78
80 | SBOM Creation | Cybersecurity & Data Privacy |  | 2025-12-15 | 2025-12-22 | 7 | 79
81 | Cybersecurity Documentation (FDA) | Cybersecurity & Data Privacy |  | 2025-12-22 | 2026-01-05 | 14 | 80
82 | Regulatory Strategy & Predicate Selection | Regulatory Submission & Documen |  | 2025-10-20 | 2025-10-27 | 7
83 | 510(k) Content Outline & Gap List | Regulatory Submission & Documen |  | 2025-10-27 | 2025-11-03 | 7 | 82
84 | Draft Device Description & Indications for Use | Regulatory Submission & Documen |  | 2025-11-03 | 2025-11-10 | 7 | 83
85 | Compile Nonclinical Test Summaries | Regulatory Submission & Documen |  | 2025-11-10 | 2025-11-24 | 14 | 84
86 | Labeling/IFU Drafts | Regulatory Submission & Documen |  | 2025-11-24 | 2025-12-01 | 7 | 85, 84
87 | 510(k) Summary & eSTAR Assembly | Regulatory Submission & Documen |  | 2025-12-01 | 2025-12-15 | 14 | 85, 86
88 | Internal Submission Review & Approval | Regulatory Submission & Documen |  | 2025-12-15 | 2025-12-22 | 7 | 87
89 | Submit 510(k) | Regulatory Submission & Documen |  | 2025-12-22 | 2025-12-24 | 2 | 88
90 | Manufacturing Plan & DMR Completion | Manufacturing Transfer & Proces |  | 2025-10-20 | 2025-11-03 | 14
91 | Pilot Build (Engineering Lot) | Manufacturing Transfer & Proces |  | 2025-11-03 | 2025-12-01 | 28 | 90
92 | Define Process Flow & Work Instructions | Manufacturing Transfer & Proces |  | 2025-12-01 | 2025-12-15 | 14 | 91, 90
93 | Tooling/Equipment IQ | Manufacturing Transfer & Proces |  | 2025-12-15 | 2025-12-29 | 14 | 92
94 | Process OQ | Manufacturing Transfer & Proces |  | 2025-12-29 | 2026-01-12 | 14 | 93
95 | Process PQ (3 lots) | Manufacturing Transfer & Proces |  | 2026-01-12 | 2026-02-09 | 28 | 94
96 | Train Production Personnel | Manufacturing Transfer & Proces |  | 2026-02-09 | 2026-02-23 | 14 | 95, 92
97 | Process Validation Report | Manufacturing Transfer & Proces |  | 2026-02-23 | 2026-03-02 | 7 | 96, 95

## Gantt
Gantt (Original)

## Gantt (Leveled)
Gantt (Leveled by Resource)

## Resource Leveling
ID | Task Name | Area (Tab) | Responsible Party | Start Date | Finish Date | Leveled Start | Leveled Finish | Resource Delay (days)
1 | Draft Design & Development Plan | Design Controls & Documentation |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
2 | Define User Needs & System Requirements | Design Controls & Documentation |  | 2025-10-27 | 2025-11-24 | 2025-10-27 | 2025-11-24 | 0
3 | Develop System Architecture & Specifications | Design Controls & Documentation |  | 2025-11-24 | 2025-12-15 | 2025-11-24 | 2025-12-15 | 0
4 | Plan Design Reviews (PDR, CDR) | Design Controls & Documentation |  | 2025-12-15 | 2025-12-22 | 2025-12-15 | 2025-12-22 | 0
5 | Create Verification Plan (trace to requirements) | Design Controls & Documentation |  | 2025-12-15 | 2025-12-29 | 2025-12-15 | 2025-12-29 | 0
6 | Generate Design Outputs (CAD, Schematics, Code baselines) | Design Controls & Documentation |  | 2025-10-20 | 2025-12-15 | 2025-10-20 | 2025-12-15 | 0
7 | Execute Design Verification (bench/system) | Design Controls & Documentation |  | 2025-12-29 | 2026-02-09 | 2025-12-29 | 2026-02-09 | 0
8 | Design Validation (simulated use) | Design Controls & Documentation |  | 2026-02-09 | 2026-03-09 | 2026-02-09 | 2026-03-09 | 0
9 | Design Transfer & DMR Compilation | Design Controls & Documentation |  | 2026-03-09 | 2026-04-06 | 2026-03-09 | 2026-04-06 | 0
10 | Create Risk Management Plan (ISO 14971) | Risk Management |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
11 | Preliminary Hazard Analysis (PHA) | Risk Management |  | 2025-10-27 | 2025-11-10 | 2025-10-27 | 2025-11-10 | 0
12 | System/Design FMEAs | Risk Management |  | 2025-11-10 | 2025-12-08 | 2025-11-10 | 2025-12-08 | 0
13 | Define Risk Controls & update requirements | Risk Management |  | 2025-12-08 | 2025-12-22 | 2025-12-08 | 2025-12-22 | 0
14 | Verify Risk Controls Implemented (tests) | Risk Management |  | 2025-12-22 | 2026-01-05 | 2025-12-22 | 2026-01-05 | 0
15 | Residual Risk Evaluation & Risk Report | Risk Management |  | 2026-01-05 | 2026-01-12 | 2026-01-05 | 2026-01-12 | 0
16 | Software Development Plan (IEC 62304) | Software Development |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
17 | Software Requirements Specification (SRS) | Software Development |  | 2025-10-27 | 2025-11-17 | 2025-10-27 | 2025-11-17 | 0
18 | Software Architecture & Detailed Design | Software Development |  | 2025-11-17 | 2025-12-08 | 2025-11-17 | 2025-12-08 | 0
19 | Implementation (Coding) | Software Development |  | 2025-12-08 | 2026-02-02 | 2025-12-08 | 2026-02-02 | 0
20 | Unit Testing | Software Development |  | 2026-02-02 | 2026-02-16 | 2026-02-02 | 2026-02-16 | 0
21 | Integration Testing (HW/SW-in-the-loop) | Software Development |  | 2026-02-16 | 2026-03-09 | 2026-02-16 | 2026-03-09 | 0
22 | Software Verification (System Tests) | Software Development |  | 2026-03-09 | 2026-04-06 | 2026-03-09 | 2026-04-06 | 0
23 | Software Validation (simulated use) | Software Development |  | 2026-04-06 | 2026-04-27 | 2026-04-06 | 2026-04-27 | 0
24 | Release & Configuration Audit | Software Development |  | 2026-04-27 | 2026-05-04 | 2026-04-27 | 2026-05-04 | 0
25 | Data Collection & Curation | AI Model Development & Validati |  | 2025-10-20 | 2025-11-10 | 2025-10-20 | 2025-11-10 | 0
26 | Dataset Split & Management | AI Model Development & Validati |  | 2025-11-10 | 2025-11-17 | 2025-11-10 | 2025-11-17 | 0
27 | Model Development & Training | AI Model Development & Validati |  | 2025-11-17 | 2025-12-22 | 2025-11-17 | 2025-12-22 | 0
28 | Internal Verification (metrics, robustness, bias) | AI Model Development & Validati |  | 2025-12-22 | 2026-01-05 | 2025-12-22 | 2026-01-05 | 0
29 | System Integration of Model | AI Model Development & Validati |  | 2025-10-20 | 2025-11-03 | 2025-10-20 | 2025-11-03 | 0
30 | Clinical/Simulated Performance Validation | AI Model Development & Validati |  | 2025-11-03 | 2025-12-01 | 2025-11-03 | 2025-12-01 | 0
31 | Algorithm Description & Documentation | AI Model Development & Validati |  | 2025-12-01 | 2025-12-08 | 2025-12-01 | 2025-12-08 | 0
32 | Schematic Design & Component Selection | Electronics Hardware Developmen |  | 2025-10-20 | 2025-11-10 | 2025-10-20 | 2025-11-10 | 0
33 | PCB Layout | Electronics Hardware Developmen |  | 2025-11-10 | 2025-12-01 | 2025-11-10 | 2025-12-01 | 0
34 | Prototype PCB Fabrication & Bring-up | Electronics Hardware Developmen |  | 2025-12-01 | 2025-12-29 | 2025-12-01 | 2025-12-29 | 0
35 | Bench/Engineering Tests | Electronics Hardware Developmen |  | 2025-12-29 | 2026-01-12 | 2025-12-29 | 2026-01-12 | 0
36 | EMC Pre-compliance Testing | Electronics Hardware Developmen |  | 2026-01-12 | 2026-01-19 | 2026-01-12 | 2026-01-19 | 0
37 | 60601 Safety Design Review | Electronics Hardware Developmen |  | 2026-01-19 | 2026-01-26 | 2026-01-19 | 2026-01-26 | 0
38 | Design Freeze (Electronics) | Electronics Hardware Developmen |  | 2026-01-26 | 2026-02-02 | 2026-01-26 | 2026-02-02 | 0
39 | Formal 60601/EMC/80601 Testing (3rd party) | Electronics Hardware Developmen |  | 2026-02-02 | 2026-02-23 | 2026-02-02 | 2026-02-23 | 0
40 | Conceptual Mechanical Design | Mechanical Design & Engineering |  | 2025-10-20 | 2025-11-10 | 2025-10-20 | 2025-11-10 | 0
41 | 3D CAD Modeling & Assemblies | Mechanical Design & Engineering |  | 2025-11-10 | 2025-12-08 | 2025-11-10 | 2025-12-08 | 0
42 | FEA & Tolerance Analysis | Mechanical Design & Engineering |  | 2025-12-08 | 2025-12-22 | 2025-12-08 | 2025-12-22 | 0
43 | Prototype Fabrication (Alpha) | Mechanical Design & Engineering |  | 2025-12-22 | 2026-01-19 | 2025-12-22 | 2026-01-19 | 0
44 | Mechanical Verification (accuracy, ROM, load) | Mechanical Design & Engineering |  | 2026-01-19 | 2026-02-09 | 2026-01-19 | 2026-02-09 | 0
45 | Design Iteration (Beta) | Mechanical Design & Engineering |  | 2025-10-20 | 2025-11-03 | 2025-10-20 | 2025-11-03 | 0
46 | Finalize Drawings & DMR | Mechanical Design & Engineering |  | 2025-11-03 | 2025-11-17 | 2025-11-03 | 2025-11-17 | 0
47 | Identify Contact Materials & Classification | Biocompatibiliity & Material Sa |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
48 | Biocompatibility Evaluation Plan (BEP) | Biocompatibiliity & Material Sa |  | 2025-10-27 | 2025-11-10 | 2025-10-27 | 2025-11-10 | 0
49 | Select GLP Test Lab & SOW | Biocompatibiliity & Material Sa |  | 2025-11-10 | 2025-11-17 | 2025-11-10 | 2025-11-17 | 0
50 | Prepare Final-finish Samples (incl. sterilization) | Biocompatibiliity & Material Sa |  | 2025-11-17 | 2025-12-01 | 2025-11-17 | 2025-12-01 | 0
51 | ISO 10993 Testing (cytotox, sensitization, irritation, etc.) | Biocompatibiliity & Material Sa |  | 2025-12-01 | 2026-01-26 | 2025-12-01 | 2026-01-26 | 0
52 | Biocompatibility Report | Biocompatibiliity & Material Sa |  | 2025-10-20 | 2025-11-03 | 2025-10-20 | 2025-11-03 | 0
53 | Select Sterilization Method & Worst-case Load | Sterilization Process Validatio |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
54 | Sterilization Validation Plan (IQ/OQ/PQ) | Sterilization Process Validatio |  | 2025-10-27 | 2025-11-10 | 2025-10-27 | 2025-11-10 | 0
55 | Contract Sterilizer Coordination (schedule) | Sterilization Process Validatio |  | 2025-11-10 | 2025-11-17 | 2025-11-10 | 2025-11-17 | 0
56 | IQ/OQ Documentation (from vendor/in-house) | Sterilization Process Validatio |  | 2025-11-17 | 2025-12-01 | 2025-11-17 | 2025-12-01 | 0
57 | PQ: Half-cycle BI Runs (3x) & Full-cycle | Sterilization Process Validatio |  | 2025-12-01 | 2025-12-29 | 2025-12-01 | 2025-12-29 | 0
58 | EO Residuals / Physical PQ & Functional Checks | Sterilization Process Validatio |  | 2025-12-29 | 2026-01-12 | 2025-12-29 | 2026-01-12 | 0
59 | Sterilization Validation Report | Sterilization Process Validatio |  | 2026-01-12 | 2026-01-19 | 2026-01-12 | 2026-01-19 | 0
60 | Package & Label Design | Packaging & Labeling Validation |  | 2025-10-20 | 2025-11-03 | 2025-10-20 | 2025-11-03 | 0
61 | Sealing Process Development | Packaging & Labeling Validation |  | 2025-11-03 | 2025-11-10 | 2025-11-03 | 2025-11-10 | 0
62 | Packaging Validation Protocol | Packaging & Labeling Validation |  | 2025-11-10 | 2025-11-17 | 2025-11-10 | 2025-11-17 | 0
63 | Build Test Samples (sterilized) | Packaging & Labeling Validation |  | 2025-11-17 | 2025-11-24 | 2025-11-17 | 2025-11-24 | 0
64 | ISTA Transport Simulation & Integrity Tests | Packaging & Labeling Validation |  | 2025-11-24 | 2025-12-15 | 2025-11-24 | 2025-12-15 | 0
65 | Accelerated Aging (if shelf-life claimed) | Packaging & Labeling Validation |  | 2025-12-15 | 2026-03-09 | 2025-12-15 | 2026-03-09 | 0
66 | Label Legibility/Adhesion Testing | Packaging & Labeling Validation |  | 2026-03-09 | 2026-03-23 | 2026-03-09 | 2026-03-23 | 0
67 | Packaging Validation Report | Packaging & Labeling Validation |  | 2026-03-23 | 2026-03-30 | 2026-03-23 | 2026-03-30 | 0
68 | Use Specification & User Profiles | Usability & Human Factors Engin |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
69 | Task Analysis & Critical Tasks | Usability & Human Factors Engin |  | 2025-10-27 | 2025-11-03 | 2025-10-27 | 2025-11-03 | 0
70 | UI/Workflow Prototyping | Usability & Human Factors Engin |  | 2025-11-03 | 2025-11-17 | 2025-11-03 | 2025-11-17 | 0
71 | Formative Usability Studies | Usability & Human Factors Engin |  | 2025-11-17 | 2025-12-01 | 2025-11-17 | 2025-12-01 | 0
72 | Summative Usability Test Protocol | Usability & Human Factors Engin |  | 2025-12-01 | 2025-12-15 | 2025-12-01 | 2025-12-15 | 0
73 | Summative Usability Test Execution | Usability & Human Factors Engin |  | 2025-12-15 | 2025-12-29 | 2025-12-15 | 2025-12-29 | 0
74 | Usability Validation Report | Usability & Human Factors Engin |  | 2025-12-29 | 2026-01-05 | 2025-12-29 | 2026-01-05 | 0
75 | Threat Modeling & Cyber Risk Assessment | Cybersecurity & Data Privacy |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
76 | Security Requirements (auth, encryption, SBOM) | Cybersecurity & Data Privacy |  | 2025-10-27 | 2025-11-03 | 2025-10-27 | 2025-11-03 | 0
77 | Implement Security Controls | Cybersecurity & Data Privacy |  | 2025-10-20 | 2025-11-17 | 2025-10-20 | 2025-11-17 | 0
78 | Vulnerability Scans & Pen Testing | Cybersecurity & Data Privacy |  | 2025-11-17 | 2025-12-01 | 2025-11-17 | 2025-12-01 | 0
79 | Remediation & Hardening | Cybersecurity & Data Privacy |  | 2025-12-01 | 2025-12-15 | 2025-12-01 | 2025-12-15 | 0
80 | SBOM Creation | Cybersecurity & Data Privacy |  | 2025-12-15 | 2025-12-22 | 2025-12-15 | 2025-12-22 | 0
81 | Cybersecurity Documentation (FDA) | Cybersecurity & Data Privacy |  | 2025-12-22 | 2026-01-05 | 2025-12-22 | 2026-01-05 | 0
82 | Regulatory Strategy & Predicate Selection | Regulatory Submission & Documen |  | 2025-10-20 | 2025-10-27 | 2025-10-20 | 2025-10-27 | 0
83 | 510(k) Content Outline & Gap List | Regulatory Submission & Documen |  | 2025-10-27 | 2025-11-03 | 2025-10-27 | 2025-11-03 | 0
84 | Draft Device Description & Indications for Use | Regulatory Submission & Documen |  | 2025-11-03 | 2025-11-10 | 2025-11-03 | 2025-11-10 | 0
85 | Compile Nonclinical Test Summaries | Regulatory Submission & Documen |  | 2025-11-10 | 2025-11-24 | 2025-11-10 | 2025-11-24 | 0
86 | Labeling/IFU Drafts | Regulatory Submission & Documen |  | 2025-11-24 | 2025-12-01 | 2025-11-24 | 2025-12-01 | 0
87 | 510(k) Summary & eSTAR Assembly | Regulatory Submission & Documen |  | 2025-12-01 | 2025-12-15 | 2025-12-01 | 2025-12-15 | 0
88 | Internal Submission Review & Approval | Regulatory Submission & Documen |  | 2025-12-15 | 2025-12-22 | 2025-12-15 | 2025-12-22 | 0
89 | Submit 510(k) | Regulatory Submission & Documen |  | 2025-12-22 | 2025-12-24 | 2025-12-22 | 2025-12-24 | 0
90 | Manufacturing Plan & DMR Completion | Manufacturing Transfer & Proces |  | 2025-10-20 | 2025-11-03 | 2025-10-20 | 2025-11-03 | 0
91 | Pilot Build (Engineering Lot) | Manufacturing Transfer & Proces |  | 2025-11-03 | 2025-12-01 | 2025-11-03 | 2025-12-01 | 0
92 | Define Process Flow & Work Instructions | Manufacturing Transfer & Proces |  | 2025-12-01 | 2025-12-15 | 2025-12-01 | 2025-12-15 | 0
93 | Tooling/Equipment IQ | Manufacturing Transfer & Proces |  | 2025-12-15 | 2025-12-29 | 2025-12-15 | 2025-12-29 | 0
94 | Process OQ | Manufacturing Transfer & Proces |  | 2025-12-29 | 2026-01-12 | 2025-12-29 | 2026-01-12 | 0
95 | Process PQ (3 lots) | Manufacturing Transfer & Proces |  | 2026-01-12 | 2026-02-09 | 2026-01-12 | 2026-02-09 | 0
96 | Train Production Personnel | Manufacturing Transfer & Proces |  | 2026-02-09 | 2026-02-23 | 2026-02-09 | 2026-02-23 | 0
97 | Process Validation Report | Manufacturing Transfer & Proces |  | 2026-02-23 | 2026-03-02 | 2026-02-23 | 2026-03-02 | 0